## Page_10
4 | GENERAL REQUIREMENTS |  | P
| c) When no specific technical requirements provided manufacturer has determined HAZARDS or HAZARDOUS SITUATIONS exists. |  | N/A
| - HAZARDS or HAZARDOUS SITUATIONS have been evaluated using the RISK MANAGEMENT PROCESS. |  | N/A
4.2.3.2 | MANUFACTURER has addressed HAZARDS or HAZARDOUS SITUATIONS not specifically addressed in the IEC 60601-1 series. |  | N/A
4.3 | Performance of clinical functions necessary to achieve iNTENDED USE or that could affect the safety of the ME EQUIPMENT or ME SYSTEM were identified during RISK ANALYSIS. | No essential performance is specified by manufacturer | N/A
| - Performance limits were identified in both NORMAL CONDITION and SINGLE FAULT CONDITION. |  | N/A
| - Loss or degradation of performance beyond the limits specified by the MANUFACTURER were evaluated |  | N/A
| - Functions with unacceptable risks are identified as ESSENTIAL PERFORMANCE………………………: |  | N/A
| - RISK CONTROL measures implemented |  | N/A
| - Methods used to verify the effectiveness of RISK CONTROL measures implemented |  | N/A
4.5 | Alternative RISK CONTROL methods utilized: |  | N/A

## Page_11
| RESIDUAL RISK resulting from the alternative RISK CONTROL measures or tests is acceptable and comparable to RESIDUAL RISK resulting from application of this standard…………………….: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | Not used alternative methods | N/A
| Alternative means based scientific data or clinical opinion or comparative studies ………………….: |  | N/A
4.6 | RISK MANAGEMENT PROCESS identifies parts that can come into contact with PATIENT but not defined as APPLIED PARTS, subjected to the requirements for APPLIED PARTS, except for Clause 7.2.10 ……………………………………..: | No such parts | N/A
| MANUFACTURER assesses the risk of accessible parts coming into contact with the patient: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
| Assessment identified the APPLIED PART TYPE requirements ……………………………………..: |  | N/A
| Components and wiring exception in the standard or by RISK MANAGEMENT PROCESS |  | N/A
| RISK MANAGEMENT PROCESS assesses components to identify components where the failure results in a HAZARDOUS SITUATION for components used outside their ratings ………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | Not used outside their ratings | N/A
| MANUFACTURER identified components where the failure results in a HAZARDOUS SITUATION ……..: |  | N/A

## Page_12
| Components determined to be acceptable where used as a MEANS OF PROTECTION ………: |  | N/A
4.9 | A COMPONENT WITH HIGH-INTEGRITY CHARACTERISTICS provided and selected appropriately……………………………………….: | Not used component with high-integrity characteristics | N/A
| RISK MANAGEMENT FILE includes an assessment to determine if the failure of components results in unacceptable RISK……………………………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
| Components identified and required to be COMPONENTS WITH HIGH INTEGRITY CHARACTERISTIC: |  | N/A
| - 250 V for HAND-HELD ME EQUIPMENT (V) ............... : |  | N/A
| – 500 V for all other ME EQUIPMENT and ME SYSTEMS |  | N/A
5.1 | Test not performed when analysis indicated condition being tested was adequately evaluated by other tests or methods ………………………..: | Not used other tests or methods | N/A

## Page_13
RISK MANAGEMENT FILE identifies combinations of simultaneous independent faults that could result in a HAZARDOUS SITUATION. (ISO 14971 Cl. 5.2-5.5) |  | N/A
c) ME EQUIPMENT with more than one RATED voltage, both a.c./ d.c. or both external power and INTERNAL ELECTRICAL POWER SOURCE tested in conditions (see 5.4) related to the least favourable voltage, nature of supply, and type of current..................................................................... : | No more than one rated voltage | N/A
d) ME EQUIPMENT intended for only d.c. supply connection tested with d.c. and influence of polarity considered ................................................. : | Not intended for only d.c. supply connection | N/A
e)ME EQUIPMENT tested with alternative ACCESSORIES and components specified in ACCOMPANYING DOCUMENTS to result in the least favourable conditions ............................................. : | No alternative accessories and components | N/A
f) ME EQUIPMENT connected to a separate power supply as specified in instructions for use | Not connected to a separate power supply | N/A

## Page_14
5.9.1 | APPLIED PARTS identified by inspection and reference to ACCOMPANYING DOCUMENTS ............... : | Dental chair, triple hot-water spray gun, triple cold-water spray gun, weak saliva pump and strong saliva pump are considered as type B applied parts | P
5.9.2.3 | Conductive parts of actuating mechanisms of electrical controls accessible after removal of handles, knobs, levers and the like regarded as ACCESSIBLE PARTS ................................................. : |  | N/A
| Conductive parts of actuating mechanisms not considered ACCESSIBLE PARTS when removal of handles, knobs, required use of a TOOL ............ : |  | N/A
| CLASS II ME EQUIPMENT, externally powered | Class I Equipment | N/A
| INTERNALLY POWERED ME EQUIPMENT |  | N/A
| EQUIPMENT with means of connection to a SUPPLY MAINS complied with CLASS I or CLASS II ME EQUIPMENT requirements when so connected, and when not connected to SUPPLY MAINS with INTERNALLY POWERED ME EQUIPMENT requirements |  | N/A
| TYPE BF APPLIED PART |  | N/A
| TYPE CF APPLIED PART |  | N/A
| DEFIBRILLATION-PROOF APPLIED PARTS |  | N/A
6.4 | ME EQUIPMENT or its parts intended to be sterilized classified according to method(s) of sterilization in instructions for use .......................... : | Not intended to be sterilized | N/A
6.5 | ME EQUIPMENT and ME SYSTEMS intended for use in an OXYGEN RICH ENVIRONMENT classified for such use and complied with 11.2.2 | Not used in oxygen rich environment | N/A

## Page_15
6.6 | CONTINUOUS or Non-CONTINUOUS OPERATION ........ : | Non-continuous operation: On: 2 min. / OFF: 18 min. | P
| Remaining markings fully recorded in ACCOMPANYING DOCUMENTS ……………………: |  | N/A
| Markings applied to individual packaging when impractical to apply to ME EQUIPMENT |  | N/A
| Single use item marked…………………………..: | No parts for single use | N/A
| Detachable components of the ME EQUIPMENT not marked; misidentification does not present an unacceptable risk, or | No detachable components | N/A
| RISK MANAGEMENT FILE includes an assessment of the RISKS relating to misidentification of all detachable parts…………………………………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.3) |  | N/A
| Detachable components of the ME EQUIPMENT are marked with the name or trademark of the MANUFACTURER, and |  | N/A
| – a MODEL OR TYPE REFERENCE |  | N/A

## Page_16
7.2.3 | Symbol 11 on Table D.1 used, optionally, advice to OPERATOR to consult ACCOMPANYING DOCUMENTS |  | N/A
7.2.4 | ACCESSORIES marked with name or trademark and contact information of their MANUFACTURER, and ………………………………………………….: | No accessories | N/A
| - with a MODEL or TYPE REFERENCE |  | N/A
| – a serial number or lot or batch identifier |  | N/A
| – the date of manufacture or use by date |  | N/A
| Markings applied to individual packaging when not practical to apply to ACCESSORIES |  | N/A
7.2.5 | ME EQUIPMENT and ME SYSTEM intended to receive power from other equipment, provided with one of the following | Not intended to receive power from other equipment | N/A
| - the name or trademark of the manufacturer of the other electrical equipment and type reference marked adjacent to the relevant connection point; or |  | N/A
| – Table D.2, SAFETY SIGN No. 10 adjacent to the relevant connection point and listing of the required details in the instructions for use; or |  | N/A
| – Special connector style used that is not commonly available on the market and listing of the required details in the instructions for use. |  | N/A
| For PERMANENTLY INSTALLED ME EQUIPMENT, NOMINAL supply voltage or range marked inside or outside of ME EQUIPMENT | Not permanently installed ME Equipment | N/A
| Multiple RATED supply voltages or multiple RATED supply voltage ranges are separated by (V/V)…..: |  | N/A
| Symbols 1-5, Table D.1 (used for same parameters………………………………………….: |  | N/A

## Page_17
| – RATED supply frequency or RATED frequency range in hertz……………………………………….: | 50/60 Hz | P
| – Symbol 9 of Table D.1 used for CLASS II ME EQUIPMENT…………………………………………..: | Class I ME Equipment | N/A
| RATED input in amps or volt-amps, or in watts when power factor exceeds 0.9 (A, VA, W)……..: |  | N/A
| RATED input for one or more RATED voltage ranges provided for upper and lower limits of the range or ranges when the range(s) is/are greater than ± 10 % of the mean value of specified range (A, VA,W)…………………………………………..: |  | N/A
| Input at mean value of range marked when range limits do not differ by more than 10 % from mean value (A, VA, W)……………………………………: |  | N/A
| Marking includes long-time and most relevant momentary volt-ampere ratings when provided, each plainly identified and indicated in ACCOMPANYING DOCUMENTS (VA)………………….: |  | N/A
| Marked input of ME EQUIPMENT provided with means for connection of supply conductors of other electrical equipment includes RATED and marked output of such means (A, VA, W)……….: |  | N/A
7.2.8 | Output connectors |  | N/A
7.2.8.2 | Output connectors are marked, except for MULTIPLE SOCKET-OUTLETS or connectors intended for specified ACCESSORIES or equipment | No electrical output connector | N/A
| TYPE BF APPLIED PARTS with symbol 20 of Table D.1: |  | N/A
| TYPE CF APPLIED PARTS with symbol 21 of Table D.1…………………………………………………..: |  | N/A

## Page_18
| DEFIBRILLATION-PROOF APPLIED PARTS marked with symbols 25-27 of Table D.1…………………: | No defibrillation-proof applied parts | N/A
| SAFETY SIGN 2 of Table D.2 placed near relevant outlet………………………………………………..: |  | N/A
| An explanation indicating protection of ME EQUIPMENT against effects of discharge of a cardiac defibrillator depends on use of proper cables included in instructions for use……………: |  | N/A
7.2.11 | ME EQUIPMENT suitable for CONTINUOUS OPERATION |  | N/A
7.2.13 | Physiological effects – SAFETY SIGN and warning statements ………………………………………..: |  | N/A
| Nature of HAZARD and precautions for avoiding or minimizing the associated RISK described in instructions for use…………………………………: (ISO 14971 Cl. 5.2-5.5, 6, 7.2) |  | N/A
7.2.14 | HIGH VOLTAGE TERMINAL DEVICES on the outside of ME EQUIPMENT accessible without the use of a TOOL marked with symbol 24 of Table D.1 | No high voltage terminal devices on the outside | N/A
7.2.15 | Requirements for cooling provisions marked……: | No cooling provisions | N/A
| Packaging marked with a suitable SAFETY SIGN indicating premature unpacking of ME EQUIPMENT could result in an unacceptable RISK…………….: |  | N/A

## Page_19
| RISK MANAGEMENT FILE includes the assessment to determine premature unpacking of ME EQUIPMENT or its parts could result in an unacceptable RISK…………………………………: (ISO 14971 Cl. 5.2-5.5, 6, 7.2) |  | N/A
| Packaging of sterile ME EQUIPMENT or ACCESSORIES marked sterile and indicates the methods of sterilization |  | N/A
7.2.19 | Symbol 7 of Table D.1 marked on FUNCTIONAL EARTH TERMINAL…………………………………….: | No functional earth terminal | N/A
7.2.20 | Removable protective means marked to indicate the necessity for replacement when the function is no longer needed………………………………..: | No removable protective means | N/A
7.2.21 | MOBILE ME EQUIPMENT marked with its mass including its SAFE WORKING LOAD in kilograms ......... | Not mobile ME Equipment : | N/A
| A marking referring to ACCOMPANYING DOCUMENTS provided for heating elements or lamp-holders designed for heating lamps that can be changed only by SERVICE PERSONNEL using a TOOL |  | N/A
7.3.2 | Symbol 24 of Table D.1, or SAFETY SIGN No.3 of Table D.2 used to mark presence of HIGH VOLTAGE parts………………………………………: | No high voltage parts | N/A
7.3.3 | Type of battery and mode of insertion marked..: | No battery used | N/A
| An identifying marking provided referring to instructions in ACCOMPANYING DOCUMENTS for batteries intended to be changed only by SERVICE PERSONNEL using a TOOL………………………….: |  | N/A
| A warning provided indicating replacement of lithium batteries or fuel cells when incorrect replacement would result in an unacceptable RISK | : | N/A

## Page_20
| RISK MANAGEMENT FILE includes an assessment to determine the replacement of lithium batteries or fuel cells leads to an HAZARDOUS SITUATION if replaced incorrectly………………………………: (ISO 14971 Cl. 5.2-5.5, 6, 7.2) |  | N/A
| ACCOMPANYING DOCUMENTS contain a warning indicating the replacement of lithium batteries or fuel cells by inadequately trained personnel could result in a HAZARDOUS SITUATION…………: |  | N/A
7.3.6 | Symbol 7 of Table D.1 marked on FUNCTIONAL EARTH TERMINALS | No functional earth terminals | N/A
7.3.7 | Terminals for supply conductors marked adjacent to terminals…………………………………………: |  | N/A
| Terminal markings included in ACCOMPANYING DOCUMENTS when ME EQUIPMENT too small to accommodate markings |  | N/A
| Terminals exclusively for neutral supply conductor in PERMANENTLY INSTALLED ME EQUIPMENT marked with Code 1 of Table D.3 |  | N/A
| Marking for connection to a 3-phase supply, complies with IEC 60445 |  | N/A
| Markings on or adjacent to electrical connection points not applied to parts requiring removal to make connection, and remained visible after connection made |  | N/A
7.3.8 | “For supply connections, use wiring materials suitable for at least X °C” or equivalent, marked at the point of supply connections |  | N/A

## Page_21
Statement not applied to parts requiring removal to make the connection, and CLEARLY LEGIBLE after connections made | N/A
– indicated by other unambiguous means | N/A
The “on” & “off” positions of switch to control power to parts of ME EQUIPMENT, marked with symbols 12 and 13 of Table D.1 or | N/A
- marked with symbols 16 and 17 of Table D.1 or | N/A
– indicated by an adjacent indicator light, or | N/A
– indicated by other unambiguous means | N/A
Switches that brings ME EQUIPMENT into “stand- by” may be indicated by symbol 29 of Table D.1 | N/A
The “on/off” positions of push button switch with bi-stable positions marked with symbol 14 of Table D.1, and | N/A
– status indicated by adjacent indicator light | N/A
– status indicated by other unambiguous means | N/A
The “on/off” positions of push button switch with momentary on position marked with symbol 15 of Table D.1 or | N/A
– status indicated by adjacent indicator light | N/A
– status indicated by other unambiguous means | N/A
RISK MANAGEMENT FILE identifies controls where a change in setting during NORMAL USE results in an unacceptable RISK…………………………….: (ISO 14971 Cl. 5.2-5.5, 6, 7.1, 7.2) | N/A
Controls provided with an associated indicating device when change of setting of a control could result in an unacceptable RISK to PATIENT in NORMAL USE………………………………………..: | N/A
– or an indication of direction in which magnitude of the function changes | N/A

## Page_22
7.4.3 | Numeric indications of parameters on ME EQUIPMENT expressed in SI units according to ISO 80000-1 except the base quantities listed in Table 1 expressed in the indicated units | N/A
| ISO 80000-1 applied for application of SI units, their multiples, and certain other units | N/A
| Affirmative statement together with SAFETY SIGN placed in instructions for use if insufficient space on ME EQUIPMENT | N/A
| Safety notices include appropriate precautions or instructions on how to reduce RISK(S) | N/A
| SAFETY SIGNs including any supplementary text or symbols described in instructions for use | N/A
| - and in a language acceptable to the intended OPERATOR | N/A

## Page_23
7.7.2 | Insulation on conductors inside ME EQUIPMENT forming PROTECTIVE EARTH CONNECTIONS identified by green and yellow at least at terminations | P
| – POTENTIAL EQUALIZATION CONDUCTORS | N/A
| – FUNCTIONAL EARTH CONDUCTORS | N/A
7.8.1 | Red indicator lights, not flashing used only for Warning | N/A
| Yellow indicator lights, not flashing used only for Caution | N/A
| Red flashing used only for HIGH PRIORITY ALARM CONDITION, interruption of current workflow needed | N/A
| Yellow flashing used only MEDIUM PRIORITY ALARM CONDITION, re-planning of workflow needed | N/A
| Yellow or Cyan, not flashing used for LOW PRIORITY ALARM CONDITION, planning of future workflow needed. | N/A
| Other colours: Meaning other than red, yellow, cyan or green (colour, meaning)…………………: | N/A
7.8.2 | Red used only for emergency control | N/A

## Page_24
– Name or trade-name of MANUFACTURER and contact information for the RESPONSIBLE ORGANIZATION can be referred to………………….: | Refer to user manual, document No.: DTC-TCF-01, version: B6, cover page | P
When ACCOMPANYING DOCUMENTS provided electronically, USABILITY ENGINEERING PROCESS includes instructions as to what is required in hard copy or as markings on ME EQUIPMENT | Hard copy provided | N/A
– known contraindication(s) to use of ME EQUIPMENT |  | N/A
Instruction for use included the following when the PATIENT is an intended OPERATOR: | Patient is not an intended operator | N/A
– the PATIENT is an intended OPERATOR |  | N/A
– warning against servicing and maintenance while the ME EQUIPMENT is in use |  | N/A
- functions the PATIENT can safely use and, where applicable, which functions the PATIENT cannot safely use; and |  | N/A
–maintenance the PATIENT can perform |  | N/A

## Page_25
| Warning statement for CLASS I ME EQUIPMENT included |  | P
| Warning statement for ME EQUIPMENT supplied with an integral MULTIPLE SOCKET-OUTLET provided | No multiple socket-outlet | N/A
| The RESPONSIBLE ORGANIZATION is referred to this standard for the requirements applicable to ME SYSTEMS | No ME systems | N/A
7.9.2.3 | Statement on ME EQUIPMENT for connection to a separate power supply provided in instructions | Not connection to a separate power supply | N/A
7.9.2.4 | Warning statement for mains- operated ME EQUIPMENT with additional power source not automatically maintained in a fully usable condition indicating the necessity for periodic checking or replacement of power source | No additional power source | N/A
| RISK MANAGEMENT FILE assesses the RISK resulting from leakage of batteries………………: (ISO 14971 Cl. 5.2-5.5, 6, 7.2) |  | N/A
| Where the RISK is unacceptable, the IFU includes a warning to remove the battery if the ME EQUIPMENT is not likely to be used for some time . | : | N/A
| Specifications of replaceable INTERNAL ELECTRICAL POWER SOURCE when provided…….: |  | N/A
| Warning indicating ME EQUIPMENT must be connected to an appropriate power source when loss of power source would result in an unacceptable RISK…………………………………: |  | N/A
| Information provided on materials and ingredients PATIENT or OPERATOR is exposed to |  | N/A
| Restrictions specified on other equipment or NETWORK/DATA COUPLINGS, other than those forming part of an ME SYSTEM, to which a SIGNAL INPUT/OUTPUT PART may be connected |  | N/A

## Page_26
APPLIED PARTS specified |  | P
Components, ACCESSORIES or ME EQUIPMENT marked for single use, except when required by MANUFACTURER to be cleaned, disinfected, or sterilized prior to use | No parts marked for single use | N/A
Instructions provided to ensure adequate maintenance of ME EQUIPMENT containing rechargeable batteries to be maintained by anyone other than SERVICE PERSONNEL | No rechargeable batteries | N/A

## Page_27
7.9.2.14 | A list of ACCESSORIES, detachable parts, and materials for use with ME EQUIPMENT provided |  | P
| Other equipment providing power to ME SYSTEM sufficiently described | No ME systems | N/A
7.9.2.17 | Instruction for use for ME EQUIPMENT emitting radiation for medical purposes, indicate the nature, type, intensity and distribution of this radiation | No emitting radiation for medical purposes | N/A
7.9.2.18 | The instructions for use for ME EQUIPMENT or ACCESSORIES supplied sterile indicate that they have been sterilized and the method of sterilization | No parts supplied sterile | N/A
| The instructions for use indicate the necessary instructions in the event of damage to the sterile packaging, and where appropriate, details of the appropriate methods of re-sterilization |  | N/A
| -special installation requirements such as the maximum permissible apparent impedance of SUPPLY MAINS |  | N/A
| -permissible range of values of inlet pressure and flow, and the chemical composition of cooling liquid | No cooling liquid | N/A

## Page_28
-description of the means for checking the oil level in partially sealed oil filled ME EQUIPMENT or its parts | No oil | N/A
-information pertaining to ESSENTIAL PERFORMANCE and any necessary recurrent ESSENTIAL PERFORMANCE and BASIC SAFETY testing including details of the means, methods and recommended frequency |  | N/A
Technical description separable from instructions for use contains required information, as follows |  | N/A
-information required by 7.2 |  | N/A
–applicable classifications in Clause 6, warning and safety notices, and explanation of SAFETY SIGNs marked on ME EQUIPMENT | Not separable from user manual | N/A
– brief description of the ME EQUIPMENT, how the ME EQUIPMENT functions and its significant physical and performance characteristics; and |  | N/A
a unique version identifier………………………..: |  | N/A
–type and full rating of fuses used in SUPPLY MAINS external to PERMANENTLY INSTALLED ME EQUIPMENT…………………………………………: | Not permanently installed me equipment | N/A

## Page_29
| – warnings identifying nature of HAZARD when replacement of a component could result in an unacceptable RISK, and when replaceable by SERVICE PERSONNEL all information necessary to safely replace the component |  | P
8.2 | Requirements related to power sources |  | N/A
8.2.1 | Connection to a separate power source |  | N/A
| When ME EQUIPMENT specified for connection to a separate power source other than SUPPLY MAINS, separate power source considered as part of ME EQUIPMENT or combination considered as an ME SYSTEM | Not connection to a separate power source | N/A
| Tests performed with ME EQUIPMENT connected to separate power supply when one specified |  | N/A
| When a generic separate power supply specified, specification in ACCOMPANYING DOCUMENTS examined |  | N/A
8.2.2 | Connection to an external d.c. power source |  | N/A
| No HAZARDOUS SITUATION as described in 13.1 developed when a connection with wrong polarity made for ME EQUIPMENT from an external d.c. source | Not connection to an external d.c. power source | N/A
| ME EQUIPMENT connected with correct polarity maintained BASIC SAFETY and ESSENTIAL PERFORMANCE |  | N/A
| Protective devices that can be reset by anyone without a TOOL returns to NORMAL CONDITION on reset |  | N/A

## Page_30
8.3 | Classification of APPLIED PARTS | P
| a) APPLIED PART specified in ACCOMPANYING DOCUMENTS as suitable for DIRECT CARDIAC APPLICATION is TYPE CF | N/A
| b) An APPLIED PART provided with a PATIENT CONNECTION intended to deliver electrical energy or an electrophysiological signal to or from PATIENT is TYPE BF or CF APPLIED PART | N/A
| c) Limits specified in b) not applied to parts when probability of a connection to a PATIENT, directly or through body of OPERATOR, is negligible in NORMAL USE, and the OPERATOR is appropriately instructed | N/A
| Voltage to earth or to other ACCESSIBLE PARTS did not exceed 42.4 V peak a.c. or 60 V d.c. for above parts in NORMAL or single fault condition (V a.c. or d.c.)……………………………………: | N/A
| Energy did not exceed 240 VA for longer than 60 s or stored energy available did not exceed 20 J at a potential of 2 V or more (VA or J)…….: | N/A
| Limits in b) does not apply to SIP/SOP connectors and separate power supply connectors if the voltage measured is less than or equal to 60 V d.c. or 42,4 V peak a.c | N/A
| d) Voltage and energy limits specified in c) above also applied to the following: | N/A
| – internal parts touchable by test pin in Fig 8 inserted through an opening in an ENCLOSURE; and | N/A

## Page_31
| – internal parts touchable by a metal test rod with a diameter of 4 mm and a length 100 mm, inserted through any opening on top of ENCLOSURE or through any opening provided for adjustment of pre-set controls by RESPONSIBLE ORGANIZATION in NORMAL USE using a TOOL |  | N/A
| Test pin or the test rod inserted through relevant openings with minimal force of no more than 1 N |  | N/A
| Test rod inserted in every possible position through openings provided for adjustment of pre-set controls that can be adjusted in NORMAL USE, with a force of 10 N |  | N/A
| Test repeated with a TOOL specified in instructions for use |  | N/A
| Test rod freely and vertically suspended through openings on top of ENCLOSURE |  | N/A
| e) Devices used to de-energize parts when an ACCESS COVER opened without a TOOL gives access to parts at voltages above levels permitted by this Clause comply with 8.11.1 for mains isolating switches and remain effective in SINGLE FAULT CONDITION |  | N/A
| A TOOL is required when it is possible to prevent the devices from operating |  | N/A
| When voltage exceeded 60 V, calculated or measured stored charge didn’t exceed 45μC….: | Not exceeded 60 V | N/A
8.4.4 | Residual voltage of conductive parts of capacitive circuits, having become accessible after ME EQUIPMENT was de-energized after removal of ACCESS COVERS, didn’t exceed 60V or calculated stored charge didn’t exceed 45μC….: |  | N/A
| A device manually discharging capacitors used when automatic discharging was not possible and ACCESS COVERS could be removed only with aid of a TOOL |  | N/A
| Capacitor(s) and connected circuitry marked with symbol 24 of Table D.1, and manual discharging device specified in technical description…………………………………………: |  | N/A

## Page_32
8.5.1 | MEANS OF PROTECTION (MOP) |  | P
| Y1 or Y2 capacitor complying with standard IEC 60384-14 considered one MEANS OF PATIENT PROTECTION ………………………………………..: | No Y1 or Y2 capacitor | N/A
| Single Y1 capacitor used for two MEANS OF PATIENT PROTECTION when the working voltage is less than 42,4 V peak a.c. or 60 V d.c………….: |  | N/A
| Two capacitors used in series, each RATED for total WORKING VOLTAGE across the pair and have the same NOMINAL capacitance |  | N/A
| Optocouplers complying with IEC 60747-5- 5:2007, or a later edition. Considered equivalent to requirements in 8.8.2 and 8.9.3 | No optocouplers | N/A
| Measurement of Air Clearance and Creepage distance on the outside |  | N/A
| Dielectric strength test across optocoupler |  | N/A

## Page_33
| Solid insulation forming a MEANS OF OPERATOR PROTECTION complied with: | See insulation diagram and insulation table | P
| – requirements of IEC 60950-1:2005, IEC 60950-1:2005/A1:2009 and IEC 60950:2005/A2:2013 or requirements of IEC 62368-1:2018 for INSULATION CO-ORDINATION |  | N/A
| – requirements of IEC 60950-1:2005, IEC 60950-1:2005/A1:2009 and IEC 60950:2005/A2:2013 or requirements of IEC 62368-1:2018 for INSULATION CO-ORDINATION |  | N/A
| PROTECTIVE EARTH CONNECTIONS forming a MEANS OF OPERATOR PROTECTION complied with Cl. 8.6 |  | N/A
| – or with requirements and tests of IEC 60950- 1:2005, IEC 60950-1:2005/A1:2009 and IEC 60950:2005/A2:2013 or requirements of IEC 62368-1:2018 for protective earthing……………: |  | N/A
| A Y2 (IEC 60384-14) capacitor is considered one MEANS OF OPERATOR PROTECTION………….: |  | N/A
| A Y1 (IEC 60384-14 ) capacitor is considered two MEANS OF OPERATOR PROTECTION…………..: |  | N/A
| Two capacitors used in series each RATED for total WORKING VOLTAGE across the pair and have the same NOMINAL capacitance |  | N/A
| Optocouplers complying with IEC 60747-5- 5:2007, or a later edition. Considered equivalent to requirements in 8.8.2 and 8.9.3 |  | N/A
| Measurement of Air Clearance and Creepage distance on the outside |  | N/A
| Dielectric strength test across optocoupler |  | N/A
8.5.2 | Separation of PATIENT CONNECTIONS |  | N/A

## Page_34
8.5.2.1 | PATIENT CONNECTIONS of F-TYPE APPLIED PART separated from all other parts by equivalent to one MEANS OF PATIENT PROTECTION for a WORKING VOLTAGE equal to the MAX. MAINS VOLTAGE……………………………………………: | No F-type applied part | N/A
| Separation requirement not applied between multiple functions of a single F-TYPE APPLIED PART |  | N/A
| PATIENT CONNECTIONS treated as one APPLIED PART in the absence of electrical separation between PATIENT CONNECTIONS of same or another function |  | N/A
| MANUFACTURER has defined if multiple functions are to be considered as all within one APPLIED PART or as multiple APPLIED PARTS……………….: |  | N/A
| Classification as TYPE BF, CF, or DEFIBRILLATION- PROOF applied to one entire APPLIED PART |  | N/A
| LEAKAGE CURRENT tests conducted per 8.7.4….: |  | N/A
| Dielectric strength test conducted per 8.8.3……: |  | N/A
| CREEPAGE and CLEARANCES measured …………: |  | N/A
| A protective device connected between PATIENT CONNECTIONS of an F-TYPE APPLIED PART and ENCLOSURE to protect against excessive voltages did not operate below 500 V r.m.s |  | N/A
8.5.2.2 | PATIENT CONNECTIONS of a TYPE B APPLIED PART not PROTECTIVELY EARTHED are separated by one MEANS OF PATIENT PROTECTION from metal ACCESSIBLE PARTS not PROTECTIVELY EARTHED…: | Metal accessible parts are protectively earthed | N/A
| – except when metal ACCESSIBLE PART is physically close to APPLIED PART and can be regarded as a part of APPLIED PART; and |  | N/A
| – RISK that metal ACCESSIBLE PART will make contact with a source of voltage or LEAKAGE CURRENT above permitted limits is acceptably low. In this case 8.7.4.7 d) does not apply |  | N/A
| LEAKAGE CURRENT tests conducted per 8.7.4…: |  | N/A
| Dielectric strength test conducted per 8.8.3 …..: |  | N/A
| Relevant CREEPAGE and CLEARANCES measured |  | N/A
| RISK MANAGEMENT FILE includes an assessment of the RISK of metal ACCESSIBLE PARTS contacting a source of voltage or LEAKAGE CURRENT above the limits……………………………………………: (ISO 14971 Cl. 5.2-5.5, 6) |  | N/A

## Page_35
8.5.2.3 | A connector on a PATIENT lead or PATIENT cable located at the end of the lead or cable distal from PATIENT, with conductive part not separated from all PATIENT CONNECTIONS by one MEANS OF PATIENT PROTECTION for a WORKING VOLTAGE equal to MAXIMUM MAINS VOLTAGE |  | N/A
| - cannot be connected to earth or hazardous voltage while the PATIENT CONNECTIONS are in contact with PATIENT………………………………: | No patient lead or patient cable | N/A
| – conductive part of connector not separated from all PATIENT CONNECTIONS did not come into contact with a flat conductive plate of not less than 100 mm diameter |  | N/A
| – CLEARANCE between connector pins and a flat surface is at least 0.5 mm |  | N/A
| – conductive part pluggable into a mains socket protected from contacting parts at MAINS VOLTAGE by insulation with a CREEPAGE DISTANCE of at least 1.0 mm, a 1500 V dielectric strength and complying with 8.8.4.1 |  | N/A
| – required test finger did not make electrical contact with conductive part when applied against access openings with a force of 10 N, |  | N/A
| Test finger test (10 N)…………………………….: |  | N/A
| Except when RISK MANAGEMENT PROCESS includes an assessment of RISKS resulting from contact with objects other than mains sockets or flat surfaces………………………………………..: (ISO 14971 Cl. 5.2-5.5, 6) |  | N/A
| – WORKING VOLTAGE for d.c. voltages with superimposed ripple was average value when peak-to-peak ripple less than 10% of average value or peak voltage when peak-to-peak ripple exceeding 10% of average value (V)……………: |  | N/A

## Page_36
| – WORKING VOLTAGE between PATIENT CONNECTIONS of an F-TYPE APPLIED PART and ENCLOSURE was highest voltage appearing across insulation in NORMAL USE including earthing of any part of APPLIED PART (V)………..: |  | N/A
| – WORKING VOLTAGE for DEFIBRILLATION-PROOF APPLIED PARTS determined disregarding possible presence of defibrillation voltages |  | N/A
| – WORKING VOLTAGE was equal to resonance voltage in case of motors provided with capacitors between the point where a winding and a capacitor are connected together and a terminal for external conductors (V)…………….: |  | N/A
8.5.5 | DEFIBRILLATION-PROOF APPLIED PARTS |  | N/A
8.5.5.1 | Classification “DEFIBRILLATION-PROOF APPLIED PART” applied to one APPLIED PART in its entirety | No defibrillation-proof applied par | N/A
| Isolation of PATIENT CONNECTIONS of a DEFIBRILLATION-PROOF APPLIED PART from other parts of ME EQUIPMENT accomplished as follows: |  | N/A
| a) No hazardous electrical energies appear during a discharge of cardiac defibrillator ……...: |  | N/A
| b) ME EQUIPMENT complied with relevant requirements of this standard, providing BASIC SAFETY and ESSENTIAL PERFORMANCE following exposure to defibrillation voltage, and recovery time stated in ACCOMPANYING DOCUMENTS……..: |  | N/A
8.5.5.2 | Means provided to limit energy delivered to a 100 Ω load…………………………………………: |  | N/A
| Parts complying with IEC 60950-1:2005, IEC 60950-1:2005/AMD1:2009 and IEC 60950- 1:2005/AMD2:2013 or IEC 62368-1:2018 for protective earthing and serving as MEANS OF OPERATOR PROTECTION but not PATIENT PROTECTION exempted from requirements of 8.6.2 to 8.6.8 |  | N/A

## Page_37
Clamping means of PROTECTIVE EARTH TERMINAL of ME EQUIPMENT for FIXED supply conductors or POWER SUPPLY CORDS comply with 8.11.4.3, and cannot be loosened without TOOL | P
Earth pin of APPLIANCE INLET forming supply connection to ME EQUIPMENT regarded as PROTECTIVE EARTH TERMINAL | N/A
except when MANUFACTURER demonstrated in RISK MANAGEMENT FILE connection will remain reliable during EXPECTED SERVICE LIFE …………: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | N/A
b) Allowable TOUCH CURRENT and PATIENT LEAKAGE CURRENT in SINGLE FAULT CONDITION were not exceeded, when impedance of PROTECTIVE EARTH CONNECTIONS exceeded values in 8.6.4 a) and Table 8.6.4, due to limited current capability of relevant circuits…………….: | N/A
DETACHABLE POWER SUPPLY CORD specified by manufacturer or delivered with product | N/A
Coating not removed when requirements for impedance and current-carrying capacity met | N/A

## Page_38
| - applied also where interchangeable parts are PROTECTIVELY EARTHED |  | N/A
8.6.7 | Terminal for connection of a POTENTIAL EQUALIZATION CONDUCTOR |  | N/A
| – Terminal is accessible to OPERATOR with ME EQUIPMENT in any position of NORMAL USE | No potential equalization terminal | N/A
| –accidental disconnection avoided in NORMAL USE |  | N/A
| – Terminal allows conductor to be detached without a TOOL |  | N/A
| – Terminal not used for a PROTECTIVE EARTH CONNECTION |  | N/A
| – Terminal marked with symbol 8 of Table D.1 |  | N/A
| – Instructions for use contain information on function and use of POTENTIAL EQUALIZATION CONDUCTOR together with a reference to requirements of this standard |  | N/A
| POWER SUPPLY CORD does not incorporate a POTENTIAL EQUALIZATION CONDUCTOR |  | N/A
8.6.8 | FUNCTIONAL EARTH TERMINAL not used to provide a PROTECTIVE EARTH CONNECTION | No functional earth terminal | N/A
8.6.9 | Class II ME EQUIPMENT |  | N/A
| Third conductor of POWER SUPPLY CORD connected to protective earth contact of MAINS PLUG provided with CLASS II ME EQUIPMENT with isolated internal screens used as functional earth connection to the screen’s FUNCTIONAL EARTH TERMINAL, coloured green and yellow | Class I ME Equipment | N/A
| ACCOMPANYING DOCUMENTS include a statement that the third conductor in the POWER SUPPLY CORD is only a functional earth. |  | N/A
| Two MEANS OF PROTECTION provided between insulation of internal screens and all internal wiring connected to them and ACCESSIBLE PARTS |  | N/A

## Page_39
– where insulation used in conjunction with a PROTECTIVE EARTH CONNECTION, insulation short circuited only under conditions in 8.6.4 b) |  | N/A
Higher values of EARTH LEAKAGE CURRENT permitted for PERMANENTLY INSTALLED ME EQUIPMENT connected to a supply circuit supplying only this ME EQUIPMENT according to local regulations or IEC 60364-7-710…………: |  | N/A
f) LEAKAGE CURRENTS flowing in a FUNCTIONAL EARTH CONDUCTOR in a non-PERMANENTLY INSTALLED ME EQUIPMENT are 5 mA in NORMAL CONDITION, 10 mA in SINGLE FAULT CONDITION….: | No functional earth conductor | N/A

## Page_40
8.8.1 | Insulation relied on as MEANS OF PROTECTION, including REINFORCED INSULATION subjected to testing |  | P
| Insulation exempted from test (complies with clause 4.8) |  | N/A
| Insulation forming MEANS OF OPERATOR PROTECTION and complying with IEC 60950-1 for INSULATION CO-ORDINATION not tested as in 8.8 |  | N/A
| b) does not form part of an ENCLOSURE and not subject to handling or abrasion during NORMAL USE, and comprised of: |  | N/A
| – at least two layers of material, each passed the appropriate dielectric strength test .................. | : | N/A
| – or three layers of material, for which all combinations of two layers together passed the appropriate dielectric strength test ........................ | : | N/A
| Dielectric strength test for one or two layers was same as for one MEANS OF PROTECTION for SUPPLEMENTARY INSULATION |  | N/A
| Dielectric strength test for one or two layers was same as for two MEANS OF PROTECTION for REINFORCED INSULATION |  | N/A
| BASIC, SUPPLEMENTARY, and REINFORCED INSULATION required between windings of wound components separated by interleaved insulation complying with a) or b), or both, except when |  | N/A
| c) Wire with solid insulation, other than solvent based enamel, complying with a) |  | N/A
| d) Wire with multi-layer extruded or spirally wrapped insulation complying with b) and complying with Annex L |  | N/A
| e) Finished wire with spirally wrapped or multi- layer extruded insulation, complying with Annex L |  | N/A
| – BASIC INSULATION: minimum two wrapped layers or one extruded layer |  | N/A

## Page_41
– SUPPLEMENTARY INSULATION: minimum two layers, wrapped or extruded | N/A
– REINFORCED INSULATION: minimum three layers, wrapped or extruded | N/A
In d) and e), for spirally wrapped insulation with CREEPAGE DISTANCES between layers less than in Table 12 or 16 (Pollution Degree 1) depending on type of insulation, path between layers sealed as a cemented joint in 8.9.3.3 and test voltages of TYPE TESTS in L.3 equal 1.6 times of normal values | N/A
Protection against mechanical stress provided where two insulated wires or one bare and one insulated wire are in contact inside wound component, crossing at an angle between 45° and 90° and subject to winding tension…………: | N/A
Tests of Annex L not repeated since material data sheets confirm compliance…………………: | N/A
Satisfactory evidence of compliance provided by manufacturer for resistance to heat…………….: | N/A

## Page_42
| b) Parts of insulating material supporting uninsulated parts of MAINS PART subjected to ball-pressure test in a), except at 125 °C ± 2 ° C or ambient indicated in technical description ±2°C plus temperature rise determined during test of 11.1 of relevant part, if higher (°C)………: | See appended Table 8.8.4.1 | P
| Insulating material with embedded heating conductors considered as one MEANS OF PROTECTION but not two MEANS OF PROTECTION |  | N/A
| Parts of natural latex rubber aged by suspending samples freely in an oxygen cylinder containing commercial oxygen to a pressure of 2.1 MPa ± 70 kPa, with an effective capacity of at least 10 times volume of samples |  | N/A
| There were no cracks visible to naked eyes after samples kept in cylinder at 70 °C ± 2 °C for 96h, and afterwards, left at room temperature for at least 16h |  | N/A
8.9.1.15 | CREEPAGE DISTANCES and AIR CLEARANCES for DEFIBRILLATION-PROOF APPLIED PARTS are 4 mm or more to meet 8.5.5.1 |  | N/A
8.9.1.16 | Conductive coatings applied to non-metallic surfaces, do not result in flaking or peeling reducing any AIR CLEARANCE or CREEPAGE DISTANCE |  | N/A

## Page_43
8.9.2 | a) Short circuiting of each single one of CREEPAGE DISTANCES and CLEARANCES in turn did not result in a HAZARDOUS SITUATION, min CREEPAGE and CLEARANCES not applied………..: | See appended Table 8.9.2 | P
8.9.3 | Spaces filled by insulating compound |  | N/A
8.9.3.1 | Only solid insulation requirements applied where distances between conductive parts filled with insulating compound | No such parts | N/A
| Thermal cycling, humidity preconditioning, and dielectric strength tests |  | N/A
8.9.3.2 | For insulating compound forming solid insulation between conductive parts, a single sample subjected to thermal cycling PROCEDURE of 8.9.3.4 followed by humidity preconditioning per 5.7 (for 48 hours), followed by dielectric strength test (cl. 8.8.3 at 1,6 x test voltage)………………: |  | N/A
| Cracks or voids in insulating compound affecting homogeneity of material didn’t occur |  | N/A
8.9.3.3 | Where insulating compound forms a cemented joint with other insulating parts, three samples tested for reliability of joint |  | N/A
| A winding of solvent-based enamelled wire replaced for the test by a metal foil or by a few turns of bare wire placed close to cemented joint, and three samples tested as follows: |  | N/A
| – One sample subjected to thermal cycling PROCEDURE of 8.9.3.4, and immediately after the last period at highest temperature during thermal cycling followed by dielectric strength test of cl. 8.8.3 at 1.6 x the test voltage ………..: |  | N/A
| – The other two samples subjected to humidity preconditioning of 5.7, except for 48 hours only followed by a dielectric strength test of cl. 8.8.3 at 1.6 times the test voltage |  | N/A
| Force was applied between bare conductors and outside metal enclosure when measuring CREEPAGE DISTANCES and AIR CLEARANCES |  | N/A

## Page_44
| RISK MANAGEMENT FILE includes an assessment of RISKS related to unwanted movement of components………………………………………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | RMF Reference to specific RISKS: risk management file, document No.: QP-0603, version: 2.0, Annex 1 and Annex 2, item 8.10.1 (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.3) | P
| Stranded conductors are not solder-coated when secured by clamping means to prevent HAZARDOUS SITUATIONS |  | N/A
8.10.3 | Interconnecting flexible cords detachable without a TOOL used provided with means for connection to comply with requirements for metal ACCESSIBLE PARTS when a connection is loosened or broken ………………………………: |  | N/A
| Other HAND-HELD parts, if disturbance or breaking of one or more of the connections could result in a HAZARDOUS SITUATION, also comply with tests of Cl. 8.11.3 |  | N/A
8.10.7 | a) Insulating sleeve adequately secured……….: | No insulating sleeve | N/A

## Page_45
b) Sheath of a flexible cord not used as a MEANS OF PROTECTION inside ME EQUIPMENT when it is subject to mechanical or thermal stresses beyond its RATED characteristics | N/A
c) Insulated conductors of ME EQUIPMENT subject to temperatures exceeding 70 °C……………….: | N/A
PERMANENTLY INSTALLED ME EQUIPMENT connected to a poly-phase SUPPLY MAINS equipped with a device not interrupting neutral conductor, provided local installation conditions prevent voltage on neutral conductor from exceeding limits in 8.4.2 c) | N/A
PERMANENTLY INSTALLED ME EQUIPMENT provided with means to isolate its circuits electrically from the SUPPLY MAINS are capable of being locked in the off position | N/A
c) A SUPPLY MAINS switch used to comply with 8.11.1 a) complies with CREEPAGE / CLEARANCES for a MAINS TRANSIENT VOLTAGE of 4 kV…………: | N/A
e) Actuator of a SUPPLY MAINS switch used to comply with 8.11.1 a) complies with IEC 60447 | N/A

## Page_46
| i) Parts within ENCLOSURE of ME EQUIPMENT with a circuit > 42.4 V peak a.c. or 60 V d.c. that cannot be disconnected from its supply by an external switch or a plug device accessible at all times is protected against touch even after opening ENCLOSURE by an additional covering |  | N/A
| A clear warning notice is marked on outside of ME EQUIPMENT to indicate it exceeds allowable touch voltage |  | N/A
| For a part that could not be disconnected from supply by an external switch or a plug device accessible at all times, the required cover or warning notice complied with this clause |  | N/A
| Standard test finger applied |  | N/A
8.11.2 | MULTIPLE SOCKET-OUTLETS integral with ME EQUIPMENT complied with 16.2 d), second dash; and 16.9.2 | No multiple socket-outlets | N/A
| Only polyvinyl chloride insulated POWER SUPPLY CORD with appropriate temperature rating used for ME EQUIPMENT having external metal parts with a temperature > 75 °C touchable by the cord in NORMAL USE ……………………………….: |  | N/A
| For ME EQUIPMENT utilizing POWER SUPPLY CORDS and operating at currents greater than 63 A, apply the electrical regulations appropriate for the jurisdiction in which the ME EQUIPMENT is to be used. |  | N/A
8.11.3.4 | APPLIANCE COUPLERS complying with IEC 60320- 1 are considered to comply with 8.11.3.5 and 8.11.3.6 ……………………………………………: |  | N/A

## Page_47
| a) Conductors of POWER SUPPLY CORD provided with strain relief and insulation protected from abrasion at point of entry to ME EQUIPMENT or a MAINS CONNECTOR by a cord anchorage |  | P
| – metal, insulated from conductive ACCESSIBLE PARTS non-PROTECTIVELY EARTHED by a MEANS OF PROTECTION, or |  | N/A
| – metal provided with an insulating lining affixed to cord anchorage |  | N/A
| d) Screws to be operated when replacing POWER SUPPLY CORD do not serve to secure any components |  | N/A
8.11.3.6 | POWER SUPPLY CORDS protected against excessive bending at inlet opening of equipment | Stationary equipment | N/A
| Cord guard complied with test of IEC 60335- 1:2001, Clause 25.14, or |  | N/A

## Page_48
ME EQUIPMENT placed such that axis of cord guard projected at an angle of 45° with cord free from stress, and a mass equal 10 x D2 gram attached to the free end of cord (g)……………..: | N/A
Cord guard of temperature-sensitive material tested at 23 °C ± 2 °C, and flat cords bent in the plane of least resistance | N/A
Curvature of the cord radius, immediately after mass attached, was not less than 1.5 x D ……..: | N/A
Terminals of components other than terminal blocks complying with requirements of this Clause and marked accordingly used as terminals intended for external conductors | N/A
Screws and nuts clamping external conductors do not serve to secure any other component | N/A

## Page_49
8.11.4.5 | Adequate space provided inside ME EQUIPMENT designed for FIXED wiring or a rewireable POWER SUPPLY CORD to allow for connection of conductors | P
| - in at least one supply lead for other single- phase CLASS II ME EQUIPMENT…………………….: | N/A
| – neutral conductor not fused for PERMANENTLY INSTALLED ME EQUIPMENT | N/A
| – fuses or OVER-CURRENT RELEASES omitted due to provision of two MEANS OF PROTECTION between all parts within MAINS PART | N/A
| Justification for omission of fuses or OVER- CURRENT RELEASES documented…………………: | N/A

## Page_50
| RISK from contact with moving parts reduced to an acceptable level using protective measures, (access, function, shape of parts, energy, speed of motion, and benefits to PATIENT considered) | P
| – Gaps in Clause 9.2.2.2, or | N/A
| – Safe distances in Clause 9.2.2.3, or | N/A
| – GUARDS and other RISK CONTROL measures in 9.2.2.4, or | N/A
9.2.2.2 | A TRAPPING ZONE considered not to present a MECHANICAL HAZARD when gaps of TRAPPING ZONE complied with dimensions per Table 20….: | N/A
9.2.2.3 | A TRAPPING ZONE considered not to present a MECHANICAL HAZARD when distances separating OPERATOR, PATIENT, and others from TRAPPING ZONES exceeded values in ISO 13857:2008 …: | N/A
9.2.2.4 | GUARDS and other RISK CONTROL measures | N/A

## Page_51
9.2.2.4.1 | A TRAPPING ZONE do not to present a MECHANICAL HAZARD when GUARDS or other RISK CONTROL measures are of robust construction, not easy to bypass or render non-operational, and did not introduce additional unacceptable RISK………………………………………………….: | N/A
9.2.2.4.2 | FIXED GUARDS held in place by systems that can only be dismantled with a TOOL | N/A
9.2.2.4.3 | Movable GUARDS that can be opened without a TOOL remained attached when GUARD was open | N/A
| – they are associated with an interlock preventing relevant moving parts from starting to move while TRAPPING ZONE is accessible, and stops movement when the GUARD is opened, | N/A
| – absence or failure of one of their components prevents starting, and stops moving parts | N/A
| Movable GUARDS complied with any applicable tests | N/A
9.2.2.4.4 | Other RISK CONTROL designed and incorporated into to the control system stops movement and | N/A
| – SINGLE FAULT CONDITIONS have a second RISK CONTROL, or | N/A
| ME EQUIPMENT is SINGLE FAULT SAFE | N/A
| c) a second RISK CONTROL provided for SINGLE FAULT CONDITION of continuous activation system, or | N/A
| - the continuous activation system is SINGLE FAULT SAFE | N/A

## Page_52
9.2.3.1 | Controls positioned, recessed, or protected by other means so that they cannot be accidentally actuated | N/A
| - unless for the intended PATIENT, the USABILITY ENGINEERING PROCESS concludes otherwise (e.g. PATIENT with special needs), or | N/A

## Page_53
| h) Emergency stopping device provided with an actuator in red and easily distinguishable and identifiable from other controls |  | P
9.2.5 | Means provided to permit quick and safe release of PATIENT in event of breakdown of ME EQUIPMENT or failure of power supply, activation of a RISK CONTROL measure, or emergency stopping ................................................................ | : | N/A
| – and uncontrolled or unintended movement of ME EQUIPMENT that could result in an unacceptable RISK prevented |  | N/A
| – Situations where PATIENT is subjected to unacceptable RISKS due to proximity of moving parts, removal of normal exit routes, or other HAZARDS prevented |  | N/A
| – Measures provided to reduce RISK to an acceptable level when after removal of counterbalanced parts, other parts of ME EQUIPMENT can move in a hazardous way |  | N/A
| RISK MANAGEMENT FILE includes an assessment of RISKS to the PATIENT related to breakdown of the ME EQUIPMENT …………………………………: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
9.4 | Instability HAZARDS |  | N/A
9.4.1 | ME EQUIPMENT and its parts, other than FIXED, for placement on a surface did not overbalance (tip over) or move unexpectedly in NORMAL USE | Fixed ME Equipment | N/A
9.4.2 | Instability – overbalance |  | N/A
9.4.2.1 | ME EQUIPMENT or its parts did not overbalance when prepared per ACCOMPANYING DOCUMENTS, or when tested …………………………………….: |  | N/A

## Page_54
9.4.2.2 | Instability excluding transport |  | N/A
| ME EQUIPMENT or its did not overbalance when placed in different positions of NORMAL USE, ........ : |  | N/A
| A warning provided when overbalance occurred during 10° inclined plane test |  | N/A
9.4.2.3 | Instability from horizontal and vertical forces |  | N/A
| a) ME EQUIPMENT or its parts with a mass of 25kg or more, intended to be used on the floor, didn’t overbalance due to pushing, leaning against it |  | N/A
| Surfaces of ME EQUIPMENT or its parts where a RISK of overbalancing exists from pushing, etc., permanently marked with a warning of the RISK |  | N/A
| ME EQUIPMENT did not overbalance when tested according to Cl. 9.4.2.3 a) |  | N/A
| b) ME EQUIPMENT, for use on the floor or on a table, did not overbalance due to sitting or stepping |  | N/A
| ME EQUIPMENT or its parts, for use on the floor or on a table, where RISK of overbalancing exists, permanently marked with the RISK warning ......... : |  | N/A
| ME EQUIPMENT did not overbalance when tested according to Cl. 9.4.2.3b)…………………………: |  | N/A
9.4.2.4 | Castors and wheels |  | N/A
9.4.2.4.1 | Means used for transportation of MOBILE ME EQUIPMENT did not result in an unacceptable RISK when MOBILE ME EQUIPMENT moved or parked in NORMAL USE | No castors and wheels | N/A
9.4.2.4.2 | Force required to move MOBILE ME EQUIPMENT did not exceed 200 N ……………………………: |  | N/A
9.4.2.4.3 | MOBILE ME EQUIPMENT exceeding 45 kg able to pass over threshold ………………………………: |  | N/A
9.4.3 | Instability from unwanted lateral movement (including sliding) |  | N/A
9.4.3.1 | a) Brakes of power-driven MOBILE ME EQUIPMENT normally activated and could only be released by continuous actuation of a control | Not mobile ME Equipment | N/A
| b) MOBILE ME EQUIPMENT provided with locking means to prevent unwanted movements |  | N/A
| c) No unwanted lateral movement resulted when MOBILE ME EQUIPMENT placed in its transport position when test per 9.4.3.1 |  | N/A
9.4.3.2 | Instability excluding transport |  | N/A

## Page_55
| a) MOBILE ME EQUIPMENT provided with wheel locks or braking system compliant with 5 tilt test ............................................................................. : |  | N/A
| b) MOBILE ME EQUIPMENT provided with wheel locks or braking system compliant with lateral stability test |  | N/A
9.4.4 | Grips and other handling devices |  | N/A
| a) ME EQUIPMENT with a mass of over 20 kg requiring lifting in NORMAL USE or transport provided with suitable handling means, or ACCOMPANYING DOCUMENTS specify safe lifting method | No grips and handling devices | N/A
| Handles, suitably placed to enable ME EQUIPMENT or its part to be carried by two or more persons and by examination of EQUIPMENT, its part, or ACCOMPANYING DOCUMENTS |  | N/A
| b) PORTABLE ME EQUIPMENT with a mass > 20 kg provided with one or more carrying-handles suitably placed to enable carrying by two or more persons as confirmed by actual carrying |  | N/A
| c) Carrying handles and grips and their means of attachment withstood loading test ................... : |  | N/A
9.5.2 | Cathode Ray tube(s) complied with IEC 60065:2001, Clause 18, or IEC 61965 ................ : | No Cathode Ray tube | N/A
| If necessary, confirmed in RISK MANAGEMENT FILE including audibility of auditory alarm signals, and PATIENT sensitivity ................................................ : |  | N/A
| If necessary, confirmed in RISK MANAGEMENT FILE including audibility of auditory alarm signals, PATIENT sensitivity, and (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A

## Page_56
| All identified RISKS mitigated to an acceptable level |  | N/A
9.6.2.2 | RISK MANAGEMENT FILE examined ......................... : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | No infrasound and ultrasound energy | N/A
9.6.3 | Hand-transmitted vibration |  | N/A
| Means provided to protect PATIENT and OPERATOR when hand-transmitted frequency- weighted r.m.s. acceleration generated in NORMAL USE exceeds specified values | No hand-transmitted vibration | N/A
| – 2.5 m/s2 for a cumulative time of 8 h during a 24 h period (m/s2) ................................................ : |  | N/A
| – Accelerations for different times, inversely proportional to square root of time (m/s2) ............ : |  | N/A
| – Elements of ME EQUIPMENT or an ACCESSORY, especially pipes and hoses leading to an unacceptable RISK protected against harmful external effects |  | N/A

## Page_57
| Means provided for isolation, or local depressurizing reservoirs and similar vessels, and pressure indication when above not possible | N/A
| – All elements remaining under pressure after isolation of ME EQUIPMENT or an ACCESSORY from its power supply resulting in an unacceptable RISK provided with clearly identified exhaust devices, and a warning to depressurize these elements before setting or maintenance activity | N/A
9.7.4 | MAXIMUM EQUIPMENT PRESSURE did not exceed MAXIMUM PERMISSIBLE WORKING PRESSURE for the part, except allowed for pressure relief devices in 9.7.7confirmed by inspection of THE MANUFACTURER'S data for the component, ME EQUIPMENT, and by functional tests ...................... : | N/A
9.7.6 | Pressure-control device regulating pressure in ME EQUIPMENT with pressure-relief device completed 100,000 cycles of operation under RATED load and prevented pressure from exceeding 90 % of setting of pressure-relief device in different conditions of NORMAL USE ...... : | N/A
9.7.7 | Pressure-relief device(s) used where MAXIMUM PERMISSIBLE WORKING PRESSURE could otherwise be exceeded met the following, as confirmed by MANUFACTURER’S data, ME EQUIPMENT, RISK MANAGEMENT FILE, and functional tests ................ : | N/A
| a) Connected as close as possible to pressure vessel or parts of system it is to protect | N/A
| b) Installed to be readily accessible for inspection, maintenance, and repair | N/A
| c) Could be adjusted or rendered inoperative without a TOOL | N/A
| d) With discharge opening located and directed as to not to release material towards any person | N/A
| e) With discharge opening located and directed as to not to deposit material on parts that could result in an unacceptable RISK | N/A

## Page_58
f) Adequate discharge capacity provided to ensure that pressure will not exceed MAXIMUM PERMISSIBLE EQUIPMENT PRESSURE of system it is connected to by more than 10 % when failure occurs in control of supply pressure | N/A
g) No shut-off valve provided between a pressure-relief device and parts it is to protect | N/A
h) Min number of cycles of operation 100 000, except for one-time use devices (bursting disks) | N/A
RISK MANAGEMENT FILE includes an assessment of the risks associated with the discharge opening of the pressure relief device .................. : ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | N/A
– Means of attachment of ACCESSORIES prevent possibility of incorrect attachment that could result in an unacceptable RISK | N/A
Additional instructions provided on checking adequacy of surface of structure parts will be attached to | N/A

## Page_59
9.8.2 | Support systems maintain structural integrity during EXPECTED SERVICE LIFE, and TENSILE SAFETY FACTORS are not less than in Table 21, except when an alternative method used to demonstrate structural integrity throughout EXPECTED SERVICE LIFE, or for a foot rest | Support systems require a tensile safety factor: 8 DTC-328 and DTC-329 were tested | P
| RISK MANAGEMENT FILE includes an assessment of the structural integrity of support system ......... : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
| All identified RISKS are mitigated to an acceptable level |  | N/A
| Where the equipment is not at equilibrium after 1 min, the RISK MANAGEMENT FILE includes an assessment of the test results ............................. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A

## Page_60
| Maximum mass of PATIENT included in SAFE WORKING LOAD of ME EQUIPMENT or its parts supporting or suspending PATIENTS adapted when MANUFACTURER specified applications |  | N/A
| Max allowable PATIENT mass < 135 kg marked on ME EQUIPMENT and stated in ACCOMPANYING DOCUMENTS |  | N/A
9.8.3.2 | a) Entire mass of PATIENT or OPERATOR distributed over an area of 0.1 m2 on a foot rest temporarily supporting a standing PATIENT or OPERATOR ............................................................. : |  | N/A
| Compliance confirmed by examination of ME EQUIPMENT specifications of materials and their processing, and tests ........................................... : |  | N/A
9.8.4 | Systems with MECHANICAL PROTECTIVE DEVICES |  | N/A
9.8.4.1 | a) A MECHANICAL PROTECTIVE DEVICE provided for the support system | No such mechanical protective device | N/A
| b) MECHANICAL PROTECTIVE complies with the requirements as follows: |  | N/A
| – Designed based on TOTAL LOAD |  | N/A
| – Has TENSILE SAFETY FACTORS for all parts not less than Table 21, row 7 |  | N/A
| – Activated before travel produced an unacceptable RISK |  | N/A
| – Considers Clauses 9.2.5 and 9.8.4.3 |  | N/A

## Page_61
| Compliance confirmed by examination of ME EQUIPMENT over travel calculations and evaluation plus functional tests ........................... : | N/A
9.8.4.2 | Activation of MECHANICAL PROTECTIVE DEVICE is made obvious to OPERATOR when ME EQUIPMENT can still be used after failure of suspension or actuation means and activation of a MECHANICAL PROTECTIVE DEVICE | N/A
| MECHANICAL PROTECTIVE DEVICE requires use of a TOOL to be reset or replaced | N/A
9.8.4.3 | MECHANICAL PROTECTIVE DEVICE intended to function once | N/A
| –use of ME EQUIPMENT not possible until replacement of MECHANICAL PROTECTIVE DEVICE . : | N/A
| – ACCOMPANYING DOCUMENTS provided with required information on replacement by service personal | N/A
| – ME EQUIPMENT permanently marked with SAFETY SIGN 2 of Table D. | N/A
| – Marking is adjacent to MECHANICAL PROTECTIVE DEVICE | N/A
| – Compliance confirmed by examination and following test ........................................................ : | N/A
| A chain, cable, band, spring, belt, jack screw nut, pneumatic or hydraulic hose, structural part or the like, employed to support a load, defeated by a convenient means causing maximum normal load to fall from most adverse position permitted by construction of ME EQUIPMENT | N/A
| Load included SAFE WORKING LOAD in 9.8.3.1 when system was capable of supporting a PATIENT or OPERATOR | N/A
| No evidence of damage to MECHANICAL PROTECTIVE DEVICE affecting its ability to perform its intended function | N/A

## Page_62
10 | PROTECTION AGAINST UNWANTED AND EXCESSIVE RADIATION |  | N/A
10.1 | X-Radiation |  | N/A
10.1.1 | The air kerma did not exceed 5 µGy/hat 5 cm from surface of ME EQUIPMENT ............................. : | No X-Radiation | N/A
| Annual exposure reduced taking into account the irradiated body part, national regulations, and/or international recommendations for ME EQUIPMENT that has permanent proximity to a PATIENT as part of the INTENDED USE |  | N/A
10.1.2 | RISK from unintended X-radiation from ME EQUIPMENT producing X-radiation for diagnostic and therapeutic purposes addressed application of applicable particular and collateral standards, or ......................................................................... : |  | N/A
| RISK MANAGEMENT PROCESS as indicated in RISK MANAGEMENT ILE.................................................: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
10.2 | RISK associated with alpha, beta, gamma, neutron, and other particle radiation, addressed in RISK MANAGEMENT PROCESS as shown in RISK MANAGEMENT FILE ................................................. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | No alpha, beta, gamma, neutron, and other particle radiation | N/A
10.3 | The power density of unintended microwave radiation at frequencies between 1 GHz and 100 GHz does not exceed 10 W/m2 | No microwave radiation | N/A
| Microwave radiation is propagated intentionally |  | N/A
10.4 | Relevant requirements of IEC 60825-1:2014 applied to lasers including laser diodes, laser light barriers or similar with a wavelength range of 180nm to 1 mm. | No lasers | N/A
10.5 | RISK associated with visible electromagnetic radiation other than emitted by lasers when applicable, addressed in RISK MANAGEMENT PROCESS as indicated in RISK MANAGEMENT FILE .. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | No visible electromagnetic radiation other than emitted by lasers and LEDS | N/A
10.6 | RISK associated with infrared radiation other than emitted by lasers addressed in RISK MANAGEMENT PROCESS as indicated in RISK MANAGEMENT FILE ................................................. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | No infrared radiation other than emitted by lasers and LEDS | N/A
10.7 | RISK associated with ultraviolet radiation other than emitted by lasers and LEDS addressed in RISK MANAGEMENT PROCESS as indicated in RISK MANAGEMENT FILE ................................................. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | No ultraviolet radiation other than emitted by lasers and LEDS | N/A

## Page_63
11 | PROTECTION AGAINST EXCESSIVE TEMPERATURES AND OTHER |  | P
11.1.2.1 | APPLIED PARTS (hot or cold intended to supply heat to a PATIENT comply ..................................... : | Not intended to supply heat to patient | N/A
| Clinical effects determined and documented in the RISK MANAGEMENT FILE (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
| Temperature (hot or cold) of APPLIED PARTS intended to supply heat to a PATIENT disclosed in the instructions for use |  | N/A
| Clinical effects with respect to characteristics taken or surface pressure documented in the RISK MANAGEMENT FILE (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A

## Page_64
| Analysis documented in the RISK MANAGEMENT FILE show that APPLIED PART temperatures are not affected by operation of the ME EQUIPMENT including SINGLE FAULT CONDITIONS. Measurement of APPLIED PART temperature according to 11.1.3 is not conducted .................. : |  | N/A
| Surfaces of APPLIED PARTS that are cooled below ambient temperatures evaluated in the RISK MANAGEMENT PROCESS ......................................... : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
11.1.3 | Measurements not made when engineering judgment and rationale by MANUFACTURER indicated temperature limits could not exceed, as documented in RISK MANAGEMENT FILE ............ : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | Measurement made | N/A
| Test corner not used where engineering judgment and rationale by MANUFACTURER indicated test corner will not impact measurements, as documented in RISK MANAGEMENT FILE ................................................. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | Test corner used | N/A
| e) Where thermal regulatory devices make this method inappropriate, alternative methods for measurement are justified in the RISK MANAGEMENT FILE .................................................. | : | N/A
11.2.2 | Me equipment and me systems used in conjunction with OXYGEN RICH ENVIRONMENTS |  | N/A
11.2.2.1 | RISK of fire in an OXYGEN RICH ENVIRONMENT reduced by means limiting spread of.................... | Not used in conjunction with :o xygen rich environments | N/A
| a) No sources of ignition discovered in an OXYGEN RICH ENVIRONMENT under any of the following conditions |  | N/A

## Page_65
1) when temperature of material raised to its ignition temperature |  | N/A
2) when temperatures affected solder or solder joints causing loosening, short circuiting, or other failures causing sparking or increasing material temperature to its ignition temperature |  | N/A
3) when parts affecting safety cracked or changed outer shape exposing temperatures higher than 300°C or sparks due to overheating |  | N/A
4) when temperatures of parts or components exceeded 300°C, atmosphere was 100 % oxygen, contact material solder, and fuel cotton |  | N/A
5) when sparks provided adequate energy for ignition by exceeding limits of Figs 35 to 37 (inclusive), atmosphere was 100 % oxygen, contact material solder, and fuel cotton |  | N/A
Deviations from worst case limits in 4) and 5) above based on lower oxygen concentrations or less flammable fuels justified and documented in RISK MANAGEMENT FILE………………………….: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
Alternative test in this clause did not identify existence of ignition sources at highest voltage or current, respectively……………………………: |  | N/A
A safe upper limit determined by dividing upper limit of voltage or current, respectively, with safety margin factor of three ................................. | : | N/A
b) RESIDUAL RISK of fire in an OXYGEN RICH ENVIRONMENT as determined by application of RISK MANAGEMENT PROCESS is based on following configurations, or in combination …….: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A
1) Electrical components in an OXYGEN RICH ENVIRONMENT provided with power supplies having limited energy levels lower than those considered sufficient for ignition in 11.2.2.1 a) as determined by examination, measurement or calculation of power, energy, and temperatures in NORMAL and SINGLE FAULT CONDITIONS identified in 11.2.3…………………………………: |  | N/A
2) Max oxygen concentration measured until it did not exceed 25 % in ventilated compartments with parts that can be a source of ignition only in SINGLE FAULT CONDITION and can be penetrated by oxygen due to an undetected leak (%)……: |  | N/A

## Page_66
| 3) A compartment with parts or components that can be a source of ignition only under SINGLE FAULT CONDITION separated from another compartment containing an OXYGEN RICH ENVIRONMENT by sealing all joints and holes for cables, shafts, or other purposes | N/A
| Effect of possible leaks and failures under SINGLE FAULT CONDITION that could cause ignition evaluated using a RISK ASSESSMENT to determine maintenance intervals by examination of documentation and RISK MANAGEMENT FILE………………………………….: | N/A
| 4) Fire initiated in ENCLOSURE of electrical components in a compartment with OXYGEN RICH ENVIRONMENT that can become a source of ignition only under SINGLE FAULT CONDITIONS self-extinguished rapidly and no hazardous amount of toxic gases reached PATIENT as determined by analysis of gases ……………….: | N/A
11.2.2.2 | RISK of ignition did not occur, and oxygen concentration did not exceed 25% in immediate surroundings due to location of external exhaust outlets of an OXYGEN RICH ENVIRONMENT | N/A
11.2.2.3 | Electrical connections within a compartment containing an OXYGEN RICH ENVIRONMENT under NORMAL USE did not produce sparks | N/A
| – Screw-attachments protected against loosening during use by varnishing, use of spring washers, or adequate torques | N/A
| – Soldered, crimped, and pin-and-socket connections of cables exiting ENCLOSURE include additional mechanical securing means | N/A
11.2.3 | SINGLE FAULT CONDITIONS related to OXYGEN RICH ENVIRONMENTS ME EQUIPMENT and ME SYSTEMS considered | N/A
| – Failure of a ventilation system constructed in accordance with 11.2.2.1 b) 2)…………………: | N/A
| – Failure of a barrier constructed in accordance with 11.2.2.1 b) 3)…………………………………: | N/A
| – Failure of a component creating a source of ignition (as defined in 11.2.2.1 a)………………..: | N/A
| – Failure of solid insulation or creepage and clearances providing equivalent of at least one MEANS OF PATIENT PROTECTION but less than two MEANS OF PATIENT PROTECTION that could create a source of ignition defined in 11.2.2.1 a)………: | N/A

## Page_67
| – Failure of a pneumatic component resulting in leakage of oxygen-enriched gas………………..: | N/A
11.3 | Constructional requirements for fire ENCLOSURES of ME EQUIPMENT | N/A
| ME EQUIPMENT met this clause for alternate means of compliance with selected HAZARDOUS SITUATIONS and fault conditions in 13.1.2……….: | N/A
| Constructional requirements were met, or | N/A
| - constructional requirements specifically analysed in RISK MANAGEMENT FILE …………….. : (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | N/A
| Justification, when requirement not met……….: | N/A
| a) Flammability classification of insulated wire and connectors within fire ENCLOSURE is minimum V-2, , when test in accordance with IEC 60695-11-10 or : | N/A
| insulated with PVC, TFE, PTFE, FEP, polychloroprene or polyimide as determined by examination of data on materials………………..: | N/A
| Flammability classification of printed circuit boards, and insulating material on which components are mounted is V-2, or better, based on IEC 60695-11-10 as decided by examination of materials data……………………: | N/A
| If no Certification, V tests based on IEC 60695- 11-10 conducted on 3 samples of complete parts (or sections of it), including area with min. thickness, ventilation openings | N/A
| b) Fire ENCLOSURE met following: | N/A
| 1) No openings at bottom or, as specified in Fig 39, constructed with baffles as in Fig 38, or made of perforated metal as in Table 25, or a metal screen with a mesh  2 × 2 mm centre to centre and wire diameter of at least 0.45 mm | N/A
| 2) No openings on the sides within the area included within the inclined line C in Fig 39 or made of perforated metal as in Table 25, or a metal screen with a mesh  2 × 2 mm centre to centre and wire diameter of at least 0.45 mm | N/A
| 3) ENCLOSURE, baffles, and flame barriers have adequate rigidity and are made of appropriate metal or of non-metallic materials……………….: | N/A
11.4 | ME EQUIPMENT and ME SYSTEMS intended for use with flammable anaesthetics | N/A

## Page_68
| ME EQUIPMENT, ME SYSTEMS and parts described in ACCOMPANYING DOCUMENTS for use with flammable with Annex G | Not intended for use with flammable anaesthetics | N/A
11.5 | ME EQUIPMENT and ME SYSTEMS intended for use in conjunction with flammable agents |  | N/A
| MANUFACTURER’S RISK MANAGEMENT PROCESS addresses possibility of fire and associated mitigations as confirmed by examination of RISK MANAGEMENT FILE ………………………………….: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | Not intended for use in conjunction with flammable agents | N/A
| Maximum fill level is indicated by marking on the ME EQUIPMENT and a warning or safety notice is given, no HAZARDOUS SITUATION (as specified in 13.1) or unacceptable RISK due to overflow developed when the reservoir or liquid storage chamber is filled to its maximum capacity and the TRANSPORTABLE ME EQUIPMENT is tilted through an angle of 10°, or for MOBILE ME EQUIPMENT exceeding 45 kg, is moved over a threshold as described in 9.4.2.4.3. |  | N/A

## Page_69
| ME EQUIPMENT and ME SYSTEMS handling liquids constructed that spillage does not wet parts as determined by review of the RISK MANAGEMENT FILE and test………………………………………..: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) | RMF Reference to specific RISKS: document No.: QP- 0603, version: 2.0, Annex 1 and Annex 2, item 11.6.3 (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.3) | P
11.6.8 | RISKS associated with compatibility of substances used with ME EQUIPMENT addressed in RISK MANAGEMENT PROCESS……………………: (ISO 14971 Cl. 5.3-5.5, 6, 7.1-7.4) |  | N/A

## Page_70
11.8 | Interruption and restoration of power supply did not result in a loss of BASIC SAFETY or ESSENTIAL PERFORMANCE |  | P
12.3 | MANUFACTURER implemented an ALARM SYSTEM compliant with IEC 60601-1-8:2006, IEC 60601- 1-8:2006/AMD1:2012 and IEC 60601-1- 8:2006/AMD2:2020……………………………….: |  | N/A
12.4 | Protection against hazardous output |  | N/A
12.4.1 | RISKS associated with hazardous output arising from intentional exceeding of safety limits addressed in RISK MANAGEMENT PROCESS………: (ISO 14971 Cl. 4.2-4.4, 5, 6.2-6.5) | No hazard output | N/A
12.4.2 | - need for indication associated with hazardous output addressed in RISK MANAGEMENT PROCESS ... (ISO 14971 Cl. 4.2-4.4, 5, 6.2-6.5) | . : | N/A
12.4.3 | RISKS associated with accidental selection of excessive output values for ME EQUIPMENT with a multi-purpose unit addressed in RISK MANAGEMENT PROCESS ......................................... : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
12.4.4 | RISKS associated with incorrect output addressed in RISK MANAGEMENT PROCESS ............ : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
12.4.5 | Diagnostic or therapeutic radiation |  | N/A
12.4.5.1 | Adequate provisions to protect OPERATORS, PATIENTS, other persons and sensitive devices in vicinity of unwanted or excessive radiation | No diagnostic or therapeutic radiation | N/A
| Radiation safety ensured by compliance with requirements of appropriate standards |  | N/A
12.4.5.2 | ME EQUIPMENT and ME SYSTEMS designed to produce X-radiation for diagnostic imaging purposes complied with IEC 60601-1-3 .............. : |  | N/A

## Page_71
12.4.5.3 | RISKS associated with radiotherapy addressed in RISK MANAGEMENT PROCESS as ............................. : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | N/A
12.4.5.4 | RISKS associated with ME EQUIPMENT producing diagnostic or therapeutic radiation other than diagnostic X-rays and radiotherapy addressed in RISK MANAGEMENT PROCESS as ......................... : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | N/A
12.4.6 | RISKS associated with diagnostic or therapeutic acoustic pressure addressed in RISK MANAGEMENT ........................................................ : (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | N/A

## Page_72
- or secondary circuits mounted on materials with a minimum flame rating of -V1, and | N/A
- Secondary circuits energized by less than 60 Vdc, 42.4 Vpeak in NC and SFC, and | N/A
- Secondary circuits limited to 100 VA or 6000 J in NC and SFC, and | N/A
- Wire insulation in secondary circuits of types PVC, TFE, PTFE, FEP, polychloroprene or polybromide | N/A
- or components in the circuit have HIGH INTEGRITY CHARACTERISTICS............................... : | N/A
– or parts and components completely contained within a fire ENCLOSURE complying with 11.3 as verified by review of design documentation | N/A
After tests of this Clause, settings of THERMAL CUT-OUTS and OVER-CURRENT RELEASES did not change sufficiently to affect their safety function | N/A

## Page_73
ME EQUIPMENT examined for compliance or appropriate tests such as dielectric strength of motor insulation according to 8.8.3 conducted | P
a 2) ME EQUIPMENT with heating elements RATED for non-CONTINUOUS OPERATION met tests | N/A
a 3) other ME EQUIPMENT with heating elements met test | N/A
Heating period stopped when a heating element or an intentionally weak part of a non-SELF- RESETTING THERMAL CUT-OUT ruptured, or current interrupted before THERMAL STABILITY without possibility of automatic restoration | N/A
Test repeated on a second sample when interruption was due to rupture of a heating element or an intentionally weak part | N/A
Both samples met 13.1.2, and open circuiting of a heating element or an intentionally weak part in second sample not considered a failure by itself | N/A
Operating period stopped when a non-SELF- RESETTING THERMAL CUT-OUT operated, or current interrupted without possibility of automatic restoration before THERMAL STABILITY | N/A
Test duration was equal to RATED operating time for non-CONTINUOUS OPERATION | N/A

## Page_74
1) Controls limiting temperature in NORMAL CONDITION disabled, except THERMAL CUT-OUTS | Thermostat short circuit. See appended Table 13.2 | P
2) When more than one control provided, they were disabled in turn |  | N/A
3) ME EQUIPMENT operated at RATED DUTY CYCLE until THERMAL STABILITY achieved, regardless of RATED operating time |  | N/A
To determine compliance with 13.2.9 and 13.2.10 motors in circuits running at 42.4 V peak a.c./ 60 V d.c. or less are covered with a single layer of cheesecloth which did not ignite during the test |  | N/A
b) Motor met running overload protection test of this clause when: |  | N/A
1) it is intended to be remotely or automatically controlled by a single control device with no redundant protection, or |  | N/A
2) it is likely to be subjected to CONTINUOUS OPERATION while unattended |  | N/A
Motor winding temperature determined during each steady period and maximum value did not exceed Table 27 (Insulation Class, Maximum temperature measured C) ................................... | : | N/A
Motor removed from ME EQUIPMENT and tested separately when load could not be changed in appropriate steps |  | N/A
Running overload test for motors operating at 42.4 V peak a.c./60 V d.c. or less performed only when examination and review of design indicated possibility of an overload |  | N/A
Test not conducted where electronic drive circuits maintained a substantially constant drive current |  | N/A

## Page_75
| Test not conducted based on other justifications (justification) ......................................................... | : | N/A
| c) ME EQUIPMENT with 3-phase motors operated with normal load, connected to a 3-phase SUPPLY MAINS with one phase disconnected, and periods of operation per 13.2.10 |  | N/A
13.2.13.4 | ME EQUIPMENT RATED for NON-CONTINUOUS OPERATION |  | N/A
| ME EQUIPMENT (other than HAND-HELD) operated under normal load and at RATED voltage or at upper limit of RATED voltage range until increase in temperature was  5 °C in one hour, or a protective device operated | The dental unit has to be kept switched on manually | N/A
| When a load-reducing device operated in NORMAL USE, test continued with ME EQUIPMENT running idle |  | N/A
| Motor winding temperatures did not exceed values in 13.2.10………………………………….: |  | N/A

## Page_80
15.3.3 | Impact test conducted……………………………: | See Appended Table 15.3 | P
15.3.4.2 | Sample of PORTABLE ME EQUIPMENT, ACCESSORIES and PORTABLE part with SAFE WORKING LOAD withstood stress as demonstrated by test………………………………………………: |  | N/A
| No damage resulting in an unacceptable RISK sustained |  | N/A
15.3.5 | MOBILE ME EQUIPMENT and MOBILE part with SAFE WORKING LOAD and in most adverse condition in NORMAL USE passed Rough Handling tests…….: | Not mobile ME Equipment | N/A
| No damage resulting in an unacceptable RISK sustained |  | N/A

## Page_81
15.4.1 | Incorrect connection of accessible connectors, removable without a TOOL, prevented where an unacceptable RISK exists,………………………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | RMF Reference to specific risks: document No.: QP- 0603, version: 2.0, Annex 1 and Annex 2, item 15.4.1 (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.3) | P
| a) Plugs for connection of PATIENT leads or PATIENT cables cannot be connected to outlets on same ME EQUIPMENT intended for other functions,…………………………………………: |  | N/A
| b) Medical gas connections on ME EQUIPMENT for different gases to be operated in NORMAL USE are not interchangeable inspection……………: |  | N/A
| b) THERMAL CUT-OUTS with a safety function with reset by a soldering not fitted in ME EQUIPMENT |  | N/A
| c) An additional independent non-SELF- RESETTING THERMAL CUT-OUT is provided………..: (ISO 14971 Cl. 5.2-5.5) |  | N/A
| e) Capacitors or other spark-suppression devices not connected between contacts of THERMAL CUT-OUTS |  | N/A
| - Positive temperature coefficient devices) complied with IEC 60730-1: 2010, Clauses 15, 17, J.15, and J.17 |  | N/A

## Page_82
| - SELF-RESETTING THERMAL CUT-OUTS and OVER- CURRENT RELEASES including circuits performing equivalent functions Certified according to appropriate standards……………………………: | See appended Table 8.10 | P
| - In the absence of Certification in accordance with IEC standards, SELF-RESETTING THERMAL CUT-OUTS and OVER-CURRENT RELEASES including circuits performing equivalent functions operated 200 times |  | N/A
| Manual reset THERMAL CUT-OUTS and OVER- CURRENT RELEASES Certified in accordance with appropriate IEC standards |  | N/A
| manual reset THERMAL CUT-OUTS and OVER- CURRENT RELEASES operated 10 times |  | N/A
| Thermal protective devices tested separately from ME EQUIPMENT when engineering judgment indicated test results would not be impacted |  | N/A
15.4.2.2 | Temperature settings clearly indicated when means provided to vary setting of THERMOSTATS |  | N/A
15.4.3 | Batteries |  | N/A
15.4.3.1 | Battery housings provided with ventilation…….: (ISO 14971 Cl. 5.2-5.5) | No batteries | N/A
| Battery compartments designed to prevent accidental short circuiting |  | N/A
15.4.3.2 | Means provided to prevent incorrect connection of polarity |  | N/A
| RISK MANAGEMENT FILE includes an assessment of RISKS associated with incorrect connection or replacement of batteries…………………………: (ISO 14971 Cl. 5.2-5.5) |  | N/A
15.4.3.3 | Overcharging of battery prevented by virtue of design………………………………………………: |  | N/A

## Page_83
| RISK MANAGEMENT FILE includes an assessment of RISKS associated with overcharging of batteries…………………………………………….: (ISO 14971 Cl. 5.2-5.5) | N/A
15.4.3.4 | Primary lithium batteries comply with IEC 60086-4 | N/A
| Secondary lithium batteries comply with IEC 62133 or IEC 62133-2 | N/A
15.4.3.5 | A properly RATED protective device provided within INTERNAL ELECTRICAL POWER SOURCE to protect against fire………………………………..: | N/A
| Protective device has adequate breaking capacity | N/A
| Justification for OVER-CURRENT RELEASES or FUSE exclusion is documented | N/A
| Short circuit test between the positive and negative poles of an INTERNAL ELECTRICAL POWER SOURCE between the output and protective device(s) omitted where 2 MOOPs provided, or | N/A
| Short circuit between the positive and negative poles of an INTERNAL ELECTRICAL POWER SOURCE between the output and protective device(s) does not result in any HAZARDOUS SITUATION | N/A
| An additional indicator light provided on ME EQUIPMENT with a stand-by state or a warm-up state exceeding 15 s, | N/A
| RISK MANAGEMENT FILE includes an assessment of RISKS associated with the use of indicator lights for EQUIPMENT incorporating non-luminous heaters……………………………………………..: (ISO 14971 Cl. 5.2-5.5) | N/A
| Requirement not applied to heated stylus-pens for recording purposes | N/A
| Indicator lights provided on ME EQUIPMENT to indicate an output exists | N/A
| Colours of indicator lights complied with 7.8.1 | N/A
| Charging mode visibly indicated | N/A

## Page_84
15.4.5 | RISKS associated with pre-set controls addressed in RISK MANAGEMENT PROCESS………: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) | No pre-set controls | N/A
15.4.6.2 | Stops on rotating/ movable parts of controls are of adequate mechanical strength ……………….: |  | N/A
| Torque values in Table 30 applied………………: |  | N/A
| No unexpected change of the controlled parameter when tested…………………………: |  | N/A
15.4.7.1 | a) HAND-HELD control devices of ME EQUIPMENT complied with 15.3.4.1 | No cord-connected hand-held control devices | N/A
| b) ENCLOSURE of foot operated control devices containing electrical circuits is at least IPX6……: |  | N/A
15.4.8 | Aluminium wires less than 16 mm2 in cross- sectional area are not used |  | N/A
15.4.9 | a) Oil container in PORTABLE ME EQUIPMENT allows for expansion of oil and is adequately sealed |  | N/A

## Page_85
| b) Oil containers in MOBILE ME EQUIPMENT sealed to prevent loss of oil during transport | N/A
| A pressure-release device operating during NORMAL USE is provided | N/A
| c) Partially sealed oil-filled ME EQUIPMENT and its parts provided with means for checking the oil level to detect leakage | N/A
| ME EQUIPMENT and technical description examined, and manual tests conducted to confirm compliance with above requirements | N/A
15.5.2 | Transformers operating at a frequency above 1kHz tested according to clause 8.8.3…………: | N/A
| - protective earth screens with a single turn have insulated overlap | N/A

## Page_86
| - insulation between primary and secondary windings complies with 8.8.2 |  | P
16 | ME SYSTEMS |  | N/A
16.1 | After installation or subsequent modification, ME SYSTEM didn’t result in an unacceptable RISK | No ME systems | N/A
| RISK MANAGEMENT FILE includes an assessment of RISKS associated with installation and modification of an ME SYSTEM…………………….: (ISO 14971 Cl. 5.2-5.5, 6) |  | N/A
| Only HAZARDS arising from combining various equipment to form a ME SYSTEM considered |  | N/A
| – ME SYSTEM provides the level of safety within the PATIENT ENVIRONMENT equivalent to ME EQUIPMENT complying with this standard |  | N/A
| – ME SYSTEM provides the level of safety outside PATIENT ENVIRONMENT equivalent to equipment complying with their respective IEC or ISO safety standards |  | N/A
| – tests performed in NORMAL CONDITION, except as specified |  | N/A
| – tests performed under operating conditions specified by MANUFACTURER of ME SYSTEM |  | N/A
| Safety tests previously conducted on individual equipment of ME SYSTEM according to relevant standards not repeated |  | N/A
| RISK MANAGEMENT methods used by MANUFACTURER of an ME SYSTEM reconfigurable by RESPONSIBLE ORGANIZATION or OPERATOR |  | N/A
| Non-ME EQUIPMENT used in ME SYSTEM complied with applicable IEC or ISO safety standards |  | N/A
| Equipment relying only on BASIC INSULATION for protection against electric shock not used in ME SYSTEM |  | N/A
16.2 | ACCOMPANYING DOCUMENTS of an ME SYSTEM |  | N/A
| Documents containing all data necessary for ME SYSTEM to be used as intended by MANUFACTURER including a contact address accompany ME SYSTEM or modified ME SYSTEM |  | N/A
| ACCOMPANYING DOCUMENTS regarded as a part of ME SYSTEM |  | N/A

## Page_87
a) ACCOMPANYING DOCUMENTS provided for each item of ME EQUIPMENT supplied by MANUFACTURER | N/A
b) ACCOMPANYING DOCUMENTS provided for each item of non-ME EQUIPMENT supplied by MANUFACTURER | N/A
c) the required information is provided: | N/A
– specifications, instructions for use as intended by MANUFACTURER, and a list of all items forming the ME SYSTEM | N/A
– instructions for installation, assembly, and modification of ME SYSTEM to ensure continued compliance with this standard | N/A
– instructions for cleaning and, when applicable, disinfecting and sterilizing each item of equipment or equipment part forming part of the ME SYSTEM | N/A
– additional safety measures to be applied during installation of ME SYSTEM | N/A
– identification of parts of ME SYSTEM suitable for use within the PATIENT ENVIRONMENT | N/A
– additional measures to be applied during preventive maintenance | N/A
– a warning forbidding placement of MULTIPLE SOCKET-OUTLET, when provided and it is a separate item, on the floor | N/A
– a warning indicating an additional MULTIPLE SOCKET-OUTLET or extension cord not to be connected to ME SYSTEM | N/A
– a warning to connect only items that have been specified as part of ME SYSTEM or specified as being compatible with ME SYSTEM | N/A
– maximum permissible load for any MULTIPLE SOCKET-OUTLET(S) used with ME SYSTEM | N/A
– instructions indicating MULTIPLE SOCKET- OUTLETS provided with the ME SYSTEM to be used only for supplying power to equipment intended to form part of ME SYSTEM | N/A
– an explanation indicating RISKS of connecting non-ME EQUIPMENT supplied as a part of ME SYSTEM directly to wall outlet when non-ME EQUIPMENT is intended to be supplied via a MULTIPLE SOCKET-OUTLET with a separating transformer | N/A

## Page_88
| – an explanation indicating RISKS of connecting any equipment supplied as a part of ME SYSTEM to MULTIPLE SOCKET-OUTLET |  | N/A
| – permissible environmental conditions of use for ME SYSTEM including conditions for transport and storage |  | N/A
| – instructions to OPERATOR not to, simultaneously, touch parts referred to in 16.4 and PATIENT |  | N/A
| d) the following instructions provided for use by RESPONSIBLE ORGANIZATION: |  | N/A
| – adjustment, cleaning, sterilization, and disinfection PROCEDURES |  | N/A
| – assembly of ME SYSTEMS and modifications during actual service life evaluated based on the requirements of this standard |  | N/A
16.3 | Instructions for use of ME EQUIPMENT intended to receive its power from other equipment in an ME SYSTEM, describe the other equipment to ensure compliance with these requirements |  | N/A
| Transient currents restricted to allowable levels for the specified IPS or UPS ................................. | : | N/A
| Technical description and installation instructions specify the actual transient currents where an IPS or UPS is not specified |  | N/A
16.4 | Parts of non-ME EQUIPMENT in PATIENT ENVIRONMENT subject to contact by OPERATOR during maintenance, calibration, after removal of covers, connectors operated at a voltage  voltage in 8.4.2 c) |  | N/A
16.5 | Safety measures incorporating a SEPARATION DEVICE applied when FUNCTIONAL CONNECTION between ME EQUIPMENT and other items of an ME SYSTEM or other systems can cause allowable values of LEAKAGE CURRENT to exceed |  | N/A
| SEPARATION DEVICE has dielectric strength, CREEPAGE and CLEARANCES required for one MEANS OF OPERATOR PROTECTION |  | N/A
| WORKING VOLTAGE was highest voltage across SEPARATION DEVICE during a fault condition, but not less than MAXIMUM MAINS VOLTAGE (V) ............ | : | N/A
16.6 | LEAKAGE CURRENTS |  | N/A
16.6.1 | TOUCH CURRENT in NORMAL CONDITION did not exceed 100μA ...................................................... | : | N/A

## Page_89
| TOUCH CURRENT did not exceed 500μA in event of interruption of any non-PERMANENTLY INSTALLED PROTECTIVE EARTH CONDUCTOR ............ | : | N/A
16.6.2 | Current in PROTECTIVE EARTH CONDUCTOR of MULTIPLE SOCKET-OUTLET didn’t exceed 5 mA....... | : | N/A
16.6.3 | PATIENT LEAKAGE CURRENT and total PATIENT LEAKAGE CURRENT of ME SYSTEM in NORMAL CONDITION did not exceed values .......................... | : | N/A
16.7 | ME SYSTEM complied with applicable requirements of Clause 9 ..................................... | : | N/A
16.8 | Interruption and restoration power to the ME SYSTEM or any part of the ME SYSTEM did not result in a loss of BASIC SAFETY or ESSENTIAL PERFORMANCE |  | N/A
16.9 | ME SYSTEM connections and wiring |  | N/A
16.9.1 | Incorrect connection of accessible connectors, removable without a TOOL, prevented where unacceptable RISK can result……………………: |  | N/A
| RISK MANAGEMENT FILE includes an assessment of RISKS associated with plugs for connection of PATIENT leads or cables likely to be located in the PATIENT ENVIRONMENT………………………..: (ISO 14971 Cl. 5.2-5.5, 6, 7.1-7.4) |  | N/A
| – Plugs for connection of PATIENT leads or PATIENT cables could not be connected to other outlets of the same ME SYSTEM likely to be located in PATIENT ENVIRONMENT, except when examination of connectors and interchanging them proved no unacceptable RISK results |  | N/A
| Medical gas connections on the ME SYSTEM for different gasses operated in NORMAL USE are not interchangeable |  | N/A
16.9.2 | MAINS PARTS, components and layout |  | N/A
16.9.2.1 | a) – MULTIPLE SOCKET-OUTLET only allows connection using a TOOL, or |  | N/A
| – MULTIPLE SOCKET-OUTLET is of a type that cannot accept MAINS PLUGS of any of the kinds specified in IEC/TR 60083, or |  | N/A
| – MULTIPLE SOCKET-OUTLET is supplied via a separating transformer |  | N/A
| b) – MULTIPLE SOCKET-OUTLET marked with SAFETY SIGN 2 of Table D.2 visible in NORMAL USE, and |  | N/A

## Page_90
– marked either individually or in combinations, with the maximum allowed continuous output in amperes or volt-amperes, or | N/A
– marked to indicate the equipment or equipment parts it may safely be attached to | N/A
– MULTIPLE SOCKET-OUTLET is a separate item or an integral part of ME EQUIPMENT or non-ME EQUIPMENT | N/A
c) MULTIPLE SOCKET-OUTLET complied with IEC 60884-1 and the following requirements: | N/A
– CREEPAGE and CLEARANCES complied with 8.9 | N/A
– It is CLASS I, and PROTECTIVE EARTH CONDUCTOR is connected to earthing contacts in socket-outlets | N/A
– PROTECTIVE EARTH TERMINALS and PROTECTIVE EARTH CONNECTIONS comply with 8.6: | N/A
– ENCLOSURE complied with 8.4.2 d) | N/A
– MAINS TERMINAL DEVICES and wiring complied with 8.11.4, when applicable | N/A
– RATINGS of components are not in conflict with conditions of use …………………………………: | N/A
– Electrical terminals and connectors of MULTIPLE SOCKET-OUTLETS prevent incorrect connection of accessible connectors removable without a TOOL | N/A
– POWER SUPPLY CORD complied with 8.11.3 | N/A
d) Additional requirements applied when MULTIPLE SOCKET-OUTLET combined with a separating transformer: | N/A
– Separating transformer complied with this standard or IEC 61558-2-1,………………………: | N/A
– Separating transformer is CLASS I | N/A
– Degree of protection against ingress of water specified as in IEC 60529 | N/A
– Separating transformer assembly marked according to 7.2 and 7.3 | N/A
– MULTIPLE SOCKET-OUTLET permanently connected to separating transformer, or socket- outlet of separating transformer assembly cannot accept MAINS PLUGS as identified in IEC/TR 60083 | N/A

## Page_91
16.9.2.2 | The impedance between the protective earth pin in the MAINS PLUG and any part that is PROTECTIVELY EARTHED and protected by only the SUPPLY MAINS circuit over-current release, did not exceed 200 m | N/A
| The impedance of an earth pathway protected by an additional intermediate circuit breaker or fuse rated 13A or lower, did not exceed 400 m | N/A
| Removal of any single item of equipment in ME SYSTEM will not interrupt the protective earthing of any other part without simultaneous disconnection of electrical supply to that part | N/A
| Additional PROTECTIVE EARTH CONDUCTORS can be detachable only by use of a TOOL | N/A
16.9.2.3 | Conductors connecting different items within an ME SYSTEM protected against mechanical damage | N/A

## Page_92
G.2.5 | ME EQUIPMENT or parts thereof for use with FLAMMABLE AESTHETIC MIXTURE WITH OXYGEN OR NITROUS OXIDE comply with G.4 and G.6 | N/A
| ME EQUIPMENT in G.2.4 to G.2.5 met appropriate tests of G.3-G.6 conducted after tests of 11.6.6 and 11.6.7 | N/A
G.3 | Marking, ACCOMPANYING DOCUMENTS | N/A
G.3.1 | CATEGORY APG ME EQUIPMENT prominently marked “APG” (symbol 23 in Table D.1) ............................. : | N/A
| Length of green-coloured band is  4 cm, and size of marking is as large as possible for particular case | N/A
| When above marking not possible, relevant information included in instructions for use ........... : | N/A
| Marking complied with tests and criteria of 7.1.2 and 7.1.3 | N/A
G.3.2 | CATEGORY AP ME EQUIPMENT prominently marked, with a green-coloured circle “AP” (symbol 22 in Table D.1) ............................................................. : | N/A
| Marking is as large as possible for the particular case | N/A
| When above marking not possible, the relevant information included in instructions for use ........... : | N/A
| Marking complied with tests and criteria of 7.1.2 and 7.1.3 | N/A
G.3.3 | The marking placed on major part of ME EQUIPMENT for CATEGORY AP or APG parts | N/A
G.3.4 | ACCOMPANYING DOCUMENTS contain an indication enabling the RESPONSIBLE ORGANIZATION to distinguish between CATEGORY AP and APG parts | N/A
G.3.5 | Marking clearly indicates which parts are CATEGORY AP or APG when only certain ME EQUIPMENT parts are CATEGORY AP or APG | N/A
G.4 | Common requirements for CATEGORY AP and CATEGORY APG ME EQUIPMENT | N/A
G.4.1 | a) CREEPAGE and CLEARANCES are according to Table 12 for one MEANS OF PATIENT PROTECTION | N/A
| b) Connections protected against accidental disconnection | N/A
| c) CATEGORY AP and APG not provided with a DETACHABLE POWER SUPPLY CORD, | N/A
G.4.2 | Construction details | N/A

## Page_93
| a) Opening of an ENCLOSURE protecting against penetration of gases or vapours into ME EQUIPMENT or its parts possible only with a TOOL | N/A
| b) ENCLOSURE complies with ................................. : | N/A
| – no openings on top covers of ENCLOSURE, | N/A
| – openings in side-covers prevented penetration of a solid cylindrical test rod | N/A
| – openings in base plates prevented penetration of a solid cylindrical test | N/A
| c) Short circuiting conductor(s) to a conductive part (when no explosive gasses) did not result in loss of integrity of the part, an unacceptable temperature, or any HAZARDOUS SITUATION | N/A
G.4.3 | a) Electrostatic charges prevented on CATEGORY AP and APG ME EQUIPMENT by a combination of appropriate measures | N/A
| – Use of antistatic materials with a limited electrical resistance............................................... : | N/A
| – Provision of electrically conductive paths from ME EQUIPMENT or its parts to a conductive floor, protective earth or potential equalization system, or via wheels to an antistatic floor | N/A
| b) Electrical resistance limits of aesthetic tubing, mattresses/ pads, castor tires & other antistatic material comply with ISO 2882 ............................. : | N/A
G.4.4 | Corona cannot be produced by components or parts of ME EQUIPMENT operating at more than 2000 V a.c. or 2400 V d.c. and not included in ENCLOSURES complying with G.5.4 or G.5.5 | N/A
G.5 | Requirements and tests for CATEGORY AP ME EQUIPMENT, parts and components | N/A
G.5.1 | ME EQUIPMENT, its parts or components do not ignite FLAMMABLE AESTHETIC MIXTURES WITH AIR under NORMAL USE and CONDITIONS based on compliance with G.5.2 to G.5.5 | N/A
| Alternatively, ME EQUIPMENT, its parts, and components complied with requirements of IEC 60079-0 for pressurized ENCLOSURES (IEC 60079- 2); for sand-filled ENCLOSURES, IEC 60079-5; or for oil immersed equipment, IEC 60079-6; and with this standard excluding G.5.2 to G.5.5.................. : | N/A
G.5.2 | Temperature limits ................................................ : | N/A

## Page_94
G.5.3 | ME EQUIPMENT, its parts, and components producing sparks in NORMAL USE and CONDITION complied with temperature requirements of G.5.2, and U and I occurring in their circuits, and max max complied as follows: | N/A
| Measured U ≤ U with I as in Fig. G.1 ......... : max zR zR | N/A
| Measured U ≤ U with C as in Fig. G.2 ....... : max c max | N/A
| Measured I ≤ I with U as in Fig G.1 .......... : max zR zR | N/A
| Measured I ≤ I with L and a U ≤ 24 V as max zL max max in Fig G.3 .............................................................. : | N/A
| – Combinations of currents and corresponding voltages within the limitations IzR.UzR ≤ 50 W extrapolated from Fig G.1 | N/A
| No extrapolation made for voltages above 42 V | N/A
| – Combinations of capacitances and corresponding voltages within limitations of C/2U2  1.2 mJ extrapolated from Fig G.2 | N/A
| No extrapolation made for voltages above 242V | N/A
| U determined using actual resistance R max | N/A
| – Combinations of currents and corresponding inductances within limitations L/2I2  0.3 mJ extrapolated from Fig G.3 | N/A
| No extrapolation made for inductances larger than 900 mH | N/A
| – U was the highest supply voltage occurring in max circuit under investigation with sparking contact open | N/A
| – I was the highest current flowing in circuit max under investigation with sparking contact closed | N/A
| – C and L taken as values occurring at the max max component under investigation producing sparks | N/A
| – Peak value considered when a.c. supplied | N/A
| – An equivalent circuit calculated to determine equivalent max capacitance, inductance, and equivalent U and I , either as d.c. or a.c. peak max max values in case of a complicated circuit.................. : | N/A
| Temperature measurements made according to 11.1, and U , I , R, L ,and C determined max max max max with application of Figs G.1-G.3 ........................... : | N/A
| Alternatively, compliance was verified by examination of design data .................................. : | N/A

## Page_95
G.5.4 | External ventilation with internal overpressure | N/A
| ME EQUIPMENT, its parts, and components enclosed in an ENCLOSURE with external ventilation by means of internal overpressure complied with the following requirements: | N/A
| a) FLAMMABLE AESTHETIC MIXTURES WITH AIR t removed by ventilation before EQUIPMENT energized, | N/A
| b) Overpressure inside ENCLOSURE was 75 Pa, min., in NORMAL CONDITION (Pa)............................ : | N/A
| Overpressure maintained at the site of potential ignition | N/A
| ME EQUIPMENT could be energized only after the required minimum overpressure was present long enough to ventilate the ENCLOSURE | N/A
| ME EQUIPMENT energized at will or repeatedly when overpressure was continuously present | N/A
| c) Ignition sources de-energized automatically when during operation overpressure dropped below 50 Pa (Pa) .................................................. : | N/A
| d) External surface of ENCLOSURE did not exceed 150 °C in 25 °C .................................................... : | N/A
G.5.5 | ENCLOSURES with restricted breathing | N/A
| ME EQUIPMENT, its parts, and components enclosed in an ENCLOSURE with restricted breathing complied with the following: | N/A
| a) A FLAMMABLE AESTHETIC MIXTURE WITH AIR did not form inside ENCLOSURE with restricted breathing | N/A
| b) Gasket or sealing material used to maintain tightness complied with aging test B-b of IEC 60068-2-2, Clause 15, at 70 °C ± 2 °C and 96 h .. : | N/A
| c) Gas-tightness of ENCLOSURE containing inlets for flexible cords maintained | N/A
| Cords are fitted with adequate anchorages to limit stresses as determined by test | N/A
| Overpressure not reduced below 200 Pa | N/A
| Tests waived when examination of ENCLOSURE indicated it is completely sealed or gas-tight without a doubt (100 % degree of certainty) | N/A
| Operating temperature of external surface of ENCLOSURE was  150 °C in 25 °C (°C)................ : | N/A

## Page_96
| Steady state operating temperature of ENCLOSURE also measured (°C) .............................................. : | N/A
G.6 | CATEGORY APG ME EQUIPMENT, parts and components thereof | N/A
G.6.1 | ME EQUIPMENT, its parts, and components did not ignite FLAMMABLE AESTHETIC MIXTURE WITH OXYGEN OR NITROUS OXIDE under NORMAL USE and SINGLE FAULT CONDITION | N/A
| ME EQUIPMENT, its parts, and components not complying with G.6.3 subjected to a CONTINUOUS OPERATION test | N/A
G.6.2 | Parts and components of CATEGORY APG ME EQUIPMENT operating in a FLAMMABLE AESTHETIC MIXTURE WITH OXYGEN OR NITROUS OXIDE supplied from a source isolated from earth by insulation equal to one MEANS OF PATIENT PROTECTION and from electrical parts by insulation twice the MEANS OF PATIENT PROTECTION ......................................... : | N/A
G.6.3 | Test of G.6.1 waived when the following requirements were met in NORMAL USE and under NORMAL and SINGLE FAULT CONDITIONS .................. : | N/A
| a) no sparks produced and temperatures did not exceed 90 °C, or | N/A
| b) a temperature limit of 90 °C not exceeded, sparks produced in NORMAL USE, and SINGLE FAULT CONDITIONS, except U and I occurring in their max max circuits complied with requirements, taking C max and L into consideration: max | N/A
| Measured U ≤ U with I as in Fig. G.4 ........ : max zR zR | N/A
| Measured U ≤ U with C as in Fig. G.5 ...... : max zC max | N/A
| Measured I ≤ I with U as in Fig G.4 .......... : max zR zR | N/A
| Measured I ≤ I with L and a U ≤ 24 V as max zL max max in Fig G.6 ............................................................. : | N/A
| – Extrapolation from Figs G.4, G.5, and G.6 was limited to areas indicated | N/A
| – U was the highest no-load voltage occurring max in the circuit under investigation, taking into consideration mains voltage variations as in Cl. 4.10 | N/A
| – I was the highest current flowing in the circuit max under investigation, considering MAINS VOLTAGE variations as in Cl. 4.10 | N/A
| – C and L are values occurring in relevant max max circuit | N/A

## Page_97
| – U additionally determined with actual max resistance R when equivalent resistance R in Fig G.5 was less than 8000  | N/A
| – Peak value considered when a.c. supplied | N/A
| – An equivalent circuit calculated to determine max capacitance, inductance, and U and I , max max either as d.c. or a.c. peak values in case of a complicated circuit ................................................ : | N/A
| – When energy produced in an inductance or capacitance in a circuit is limited by voltage or current-limiting devices, two independent components applied, to obtain the required limitation even when a first fault (short or open circuit) in one of these components | N/A
| - requirement not applied to transformers complying with this standard | N/A
| - requirement not applied to wire-wound current- limiting resistors provided with a protection against unwinding of the wire in case of rupture | N/A
| Compliance verified by examination of CATEGORY APG ME EQUIPMENT, parts, and components , or | N/A
| Temperature measurements made in accordance with 11.1 ............................................................... : | N/A
| - or U , I , R, L and C determined max max max max together with application of Figs G.4-G.6 ............ : | N/A
| Alternatively, compliance verified by comparison with design data ................................................... : | N/A
G.6.4 | ME EQUIPMENT, its parts, and components heating a FLAMMABLE AESTHETIC MIXTURE WITH OXYGEN OR NITROUS OXIDE provided with a non-SELF- RESETTING THERMAL CUT-OUT and complied with 15.4.2.1 ............................................................................... : | N/A
| Current-carrying part of heating element is not in direct contact with FLAMMABLE AESTHETIC MIXTURE WITH OXYGEN OR NITROUS OXIDE | N/A
G.7 | Test apparatus for flammable mixtures according to this Clause and Fig G.7 | N/A

## Page_98
L.1 | BASIC, SUPPLEMENTARY, DOUBLE, and REINFORCED INSULATION in wound components without interleaved insulation complied with this Annex | N/A
L.2 | Wire construction | N/A
| Overlap of layers when wire is insulated with two or more spirally wrapped layers of tape is adequate to ensure continued overlap during manufacture of wound component | N/A
| Layers of spirally wrapped wire insulation are sufficiently secured to maintain the overlap | N/A
L.3 | Type Test | N/A
| The wire subjected to tests of L.3.1 to L.3.4 at a temperature and a relative humidity specified | N/A
L.3.1 | Dielectric strength | N/A
| Dielectric strength test of Clause 8.8.3 for the appropriate type and number of MOP(s) conducted with no breakdown: | N/A
| – 3000 V for BASIC and SUPPLEMENTARY INSULATION (V) ...................................................... : | N/A
| – 6000 V for REINFORCED INSULATION (V) .............. : | N/A
L.3.2 | Flexibility and adherence | N/A
| Sample subjected to flexibility and adherence | N/A
| Sample examined per IEC 60851-3: 1997, cl. 5.1.1.4, followed by dielectric test of cl. 8.8.3, with no breakdown | N/A
| Test voltage was at least the voltage in Tables 6 and 7 but not less than the following: | N/A
| – 1500 V for BASIC and SUPPLEMENTARY INSULATION (V) ...................................................... : | N/A
| – 3000 V for REINFORCED INSULATION (V) ............. : | N/A
| Tension applied to wire during winding on mandrel calculated from the wire diameter equivalent to 118 MPa ± 11.8 MPa .......................................... : | N/A
L.3.3 | Heat Shock | N/A
| Sample subjected to heat shock test 9 of IEC 60851-6:1996, followed by dielectric strength test of clause 8.8.3 | N/A

## Page_99
| Test voltage was at least the voltage in Tables 6 and 7, but not less than the following: | N/A
| – 1500 V for BASIC and SUPPLEMENTARY INSULATION (V) ...................................................... : | N/A
| – 3000 V for REINFORCED INSULATION (V) ............. : | N/A
| Mandrel diameter and tension applied as in clause L.3.2, (MPa; N/mm2) ............................................ : | N/A
| Dielectric strength test conducted at room temperature after removal from the oven | N/A
L.3.4 | Retention of electric strength after bending | N/A
| Five samples prepared as in L.3.2 subjected to dielectric strength and bending tests | N/A
| Test voltage was at least the voltage in Tables 6 and 7, but not less than the following: | N/A
| – 1500 V for BASIC and SUPPLEMENTARY INSULATION (V) ...................................................... : | N/A
| – 3000 V for REINFORCED INSULATION (V) ............. : | N/A
| Test voltage applied between the shot and conductor | N/A
| Mandrel diameter and tension applied as in L.3.2, (MPa; N/mm2)....................................................... : | N/A
L.4 | Tests during manufacture | N/A
L.4.1 | Production line dielectric strength tests done by the manufacture per L.4.2 and L.4.3 .................... : | N/A
L.4.2 | Test voltage for routine testing (100 % testing) is at least the voltage in Tables 6 and 7 but not less than the following: | N/A
| – 1500 V r.m.s. or 2100 V peak for BASIC and SUPPLEMENTARY INSULATION (V) ............................ : | N/A
| – 3000 V r.m.s. or 4200 V peak for REINFORCED INSULATION (V) ...................................................... : | N/A
L.4.3 | Sampling tests conducted using twisted pair samples (IEC 60851-5:1996, clause 4.4.1) ......... : | N/A
| Minimum breakdown test voltage at least twice the voltage in Tables 6 and 7 but not less than: | N/A
| – 3000 V r.m.s. or 4200 V peak for BASIC and SUPPLEMENTARY INSULATION................................... : | N/A
| – 6000 V r.m.s. or 8400 V peak for REINFORCED INSULATION ............................................................. : | N/A

## Page_111
Type CF AP: 10 µA NC;50 µA SFC (d.c. or a.c. current)
Fig. 18 and 20 – Total Patient Leakage | — | Maximum allowed values: Type B & BF: 1000 µA Type CF: N/A